vc/20-Oct-03

### Week 42 starting 13 Oct 03

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Schedule for this week :

| BENCH | CURRENT MAGNET | STATUS Today(14/10/03) | Forward look, & Next Magnet |
| --- | --- | --- | --- |
| TBA1 | 3042 | ONLY FOR LEAK TEST | Bench de-activated – preparation for D bench |
| TBA2 | 1020 | Final Cold Tests | Next Magnet 2024? |
| TBC1 | 3045 | Disconnection | Next Magnet 3044 ? |
| TBC2 | SSS1 | cold | Cold tests |
| TBE1 | 1042 | Nearly at 1.9K (at 1400 hrs ?) | START COLD TESTS |
| TBE2 | 3039 | Disconnection | Next 1050 ? |
| TBF1 | 2002 after F2 done |  | Prepare For Commissioning AFTER F2 |
| TBF2 | 2025 | Cold | Commissioning TESTS |

Test Coordinator this week:  V Chohan;  MTM  Tech. on call:  (16-3602) : Richard

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

  - Loading up time &  AQA Analyses files(Daniel)

  - Traveller Followup – “I” shift 0830-1730

  - F Bench Commissioning

  - Interactions in Traveller & Info  ( Q-Table, HV Tests)

## New Test Sequence Update Web form & Display(GHH)

  - AOB

- Problems from Operation Last week (Stephane + Test Supervisor )

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 7_M | TBE1 : Problem with power supply. D_2_U-L pot-aim card did not work. | MTM |  | Investigation by Maryline-Richard Maybe related to cables disconnected during HV tests. |
| Tuesday 7_A | TBE2: Quench at 1.5 kA (all HF ) Triggered in ext bus bar. Symmetrical quench | MTM |  | Careful analysis in manual mode. |
| Tuesday 7_A | TBE2: Quench at 1.5 kA (all LF ) done in MM position. | MTM |  | Repeated. |
| Tuesday 7_A | TBE2: MM0 (MM position checked). Over-range appeared in 20-21 | MTM |  | Repeated and OK |
| Wednesday 8_A | TBE1: IAP @ warm Card 7 in the IAP rack (2 channels ) damaged because D20 connector not plugged properly during HV test. | MTM |  | Follow the procedure. |
| Wednesday 8_N | MM in TBE1 (shaft 24): Problem in the start point of the rotation | MTM |  | Mini-load line and tightening of shaft with TRU Investigation on progress. |
| Thursday 9_M | MM TBE2: Angle difference in aperture 2. Integrator not properly configured. | MTM |  | Solved. |
| Thursday 9_A | HV test in TBF1: Problem to launch new application HV | AB/CO | HV | Open vi and launch hv.vi |
| Friday 10_M | TBE1: IAP at warm not possible. Quench heater could not be charged. | MTM |  |  |
| Friday 10_A | LF web not starting for TBF1, 2 |  |  |  |
| Friday 10_A | IAP at warm on TBE1 not launched Mag_air_temp error. Hf files corrupted | MTM? | IAP | Done manually. |
| Friday 10_A | T coil (TBA2) cable connecting the power supply (1.5 mA) out of use. | MTM |  | Cable from TBE used. |
| Saturday 11_M | PLC interlock on current lead A (TBC1) .TE131 shows 547K | MTM/ACR |  | Correct signal simulated by Cryo. |
| Saturday 11_M | TBF2 : Connection of Shaft Cables missing and Tcoil cables for power supply as well | MTM |  |  |
| Saturday 11_M | IAP on TBA2: 3 offset could not be adjusted. 3 cards used S08_e coil which is broken | MTM |  |  |
| Saturday 11_A_M | TBC1: Quench at 1.9 K AQA gives time error (-48 s) for Q.heaters | MTM | AQA |  |
| Saturday 11_M | IAP on TBA2: 3 offset could not be adjusted. 3 cards used S08_e coil which is broken | MTM |  |  |
| Saturday 11_N | HV at cold could not start. On TBA2 server | AB/CO | HV | Run on TBA1 |
| Saturday 11_N | TBA2: Main current leads not connected | MTM |  | Test coordinator has to foreseen this. |
| Saturday 11_N | TBC1: Tcoils errors in sector 11_12 | MTM |  |  |
| Sunday 12_M | TBA2: MM position Overrange in sector 20-21 | MTM |  | Rotation too slow (20%) Good at 80 % Setting will be modified. |
| Sunday 12_A | TBA2: Quench at 1.5 kA (all HF): quench in Vbus Error in the YT111 heater discharges. Repeated : same results but error in the four Qheater discharge | MTM | AQA | Quench symmetrical? |
| Sunday 12_A | TBA2: Quench at 1.5 kA (all HF): repeated | ACR |  | Faulty trigger |
| Sunday 12_N | TBA2: Wrong current in one quench heater | MTM |  | Stop the test. Investigation by Maryline. Remove the double acquisition of LF. Origin of the parasitic signal not clear. |
| Sunday 12_N | Training quench 1 in TBA2: MIITS exceeded the limit | MTM |  |  |
| Sunday 12_N | Training quench in TBA2: LF file not found automatically in file manager. | MTM |  | Analysis done manually. Access with LF off and HF off done during the transfer processing? |
| Monday 13_M | Training quench 3 (TBA2): LF file not transferred. Copy not performed correctly. File not readable | MTM |  |  |