vc/5-Nov-03

### Week 45 starting 03 nov 03

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Schedule for this week :

| BENCH | CURRENT MAGNET | STATUS Today(04/11/03) | Forward look, & Next Magnet |
| --- | --- | --- | --- |
| TBA1 | NOT Active |  | Bench de-activated – preparation for D bench “Park” here 2001 ( moved from D bench ) |
| TBA2 | 3055 | Start to connect | Magnet 3055 should or just arrived |
| TBC1 | 1048 | Under connection |  |
| TBC2 | SSS3 | To connect |  |
| TBE1 | 2026 | Cooling after TH.Cycle | COLD Tests AFTER THERMAL CYCLE foresee on Wednesday |
| TBE2 | 3008 | Cold | Continue cold tests |
| TBF1 | 1052 | Warm; to remove | 1021 next ? |
| TBF2 |  | CFB under repair ? |  |

Test Coordinator this wek:  V.Chohan;  MTM  Tech. on call:  (16-3602) :

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

  - Resistance M’ments & Traveller

  - new todo list  numbers…

  - AOB

Current (wed 5thNov  pm) Bench Status:

SM18 LHC Magnet Test Facility

| TBA1 | _ | _ |  |
| --- | --- | --- | --- |

| TBA2 | MBAR3055 | Prep7 HV Initial @ Warm | 300 K |
| --- | --- | --- | --- |

| TBB1 | _ | _ |  |
| --- | --- | --- | --- |

| TBB2 | _ | _ |  |
| --- | --- | --- | --- |

| TBC1 | MBBL1048 | Prep12 Pumping + Cool Down | Cooling |
| --- | --- | --- | --- |

| TBC2 | SSS003 | Prep2 HV Test Warm | 300 K |
| --- | --- | --- | --- |

| TBD1 | _ | _ |  |
| --- | --- | --- | --- |

| TBD2 | _ | _ |  |
| --- | --- | --- | --- |

| TBE1 | MBAR2026 | PT6.1 Training 1 | 1.9 K |
| --- | --- | --- | --- |

| TBE2 | MBAR3008 | MM13 TRU Disconnect | Warming |
| --- | --- | --- | --- |

| TBF1 | MBAR1052 | ICS 3Traveller Last | 300 K |
| --- | --- | --- | --- |

| TBF2 | MBBR1039 | Prep3 HV Initial @ Warm | 300 K |
| --- | --- | --- | --- |

Problems of Last Week

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 28_M | TBF2: LHC cycle : motor giving problem CPU error on MMP | MTM/CO |  | LHC cycle aborted and repeated later Problem solved by Pedro?. |
| Tuesday 28_M | TBE1: HV test at warm failed on dipole-QH-ground. | MTM |  | Problem with ICS rack? |
| Tuesday 28_A | TBF2: After a Min energy quench alarm in AQA : High DT (144 V) on D2_L. Voltage sum not respected on D2 pole | MTM |  | Stephane changed the card. |
| Tuesday 28_N | Cryo problems : many stops during the nights | ACR |  |  |
| Tuesday 28_N Wednesday 29_M | TBF2: Motor problem son TRU when launching the auto-test. Encoders errors, motor busy |  |  | Origin on investigation. MMP launched by switching off/on the power supply of the crate |
| Wednesday 29_M | TBF1: IAP @warm all LF heaters. Wrong storage of the file | MTM | IAP | The correct active bench has to be selected. Active bench selected automatically when the sequence is selected? |
| Wednesday 29_A | TBC1: diode test reverse mode, Impossible to transfer the file on sun MTA19 | MTM |  | O.Berrig modified badly the .tchrc |
| Wednesday 29_N | TBA2: MM0: MUX problem during auto test | MTM |  |  |
| Thursday 30_M | TBA2: during pre-cycle (LHC) power application was hung and not communication to power converter. | AB/PO | Power | LHC cycle re-started |
| Thursday 30_N | TBA2: de-excitation test: After 9000 A the current did not raise further. | AB/PO | Power | The button execute ramp has to be clicked. |
| Friday 31_N | TBE1: IAP at cold : offset error in Vgas_inT_lead, D1_L_U, D1-D2. 8 errors in QA compensation. | MTM | IAP | Adjusted . |
| Friday 31_N | TBE1: Provoked quench at 1.5 kA (HF). Quench in Vbus ext (repeated 3 times). |  |  | Symmetrical quench? |
| Saturday 1_M | TBE1: Min energy at 3 kA (diode test) Data for diode test not in order ??? | MTM |  |  |
| Saturday 1_A | TBF1: T coil .D2S01_C opened, (c coil in shaft 25 broken), wrong value in D1_S07 | MTM |  |  |
| Saturday 1_A | TBF1: HV test cold before test: Cryo _X1 not OK Communication problem between PLC cryo and PLC | ACR/MTM |  |  |
| Saturday 1_N | Cryo problem: Difficulty to control the pressure of 3 magnets due to lack of He pressure | ACR |  |  |
| Saturday 1_N Sunday_A | TBF1: Quench at 1.5 kA and 3 kA: Signals in D2 are zero and D1-D2 negative! Non consistence between the signals | MTM |  | Test stopped. MBAL template used instead of MBBR |
| Saturday 1_N | TBE1: LHC cycle aborted because the cable were scratching with the cover. | MTM |  | Test stopped. |
| Sunday 2_A | TBA2: IAP final at warm power supply is tripping No automatic trigger | MTM | IAP | Test stopped. |
| Sunday 2_A | TBA2: QH resistance : values are found to be too high. Problem in the connectors of the measurement box Request of cables fitted to measured the cryoheater (as for the temperature sensor). One connector in the cable of the 4-wires measurement for the resistance is out of work. | MTM | IAP | Box replaced. Measurement performed again. |