vc/11-Nov-03

### Week 46 starting 10 nov 03

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Schedule for this week :

| BENCH | CURRENT MAGNET | STATUS Today(11/11/03) | Forward look, & Next Magnet |
| --- | --- | --- | --- |
| TBA1 | NOT Active |  | Bench de-activated – preparation for D bench “Park” here 2001 ( moved from D bench ) |
| TBA2 | 3055 | warming | warm@2400 hrs 11/11 2029 next ?or thermal cycle |
| TBC1 | 1048 | To disconnect | 3027 next ? |
| TBC2 | SSS3 | cold |  |
| TBE1 | 1034 | connection | Start tests after E2 finished |
| TBE2 | 3008 | Cooling? | cold@22hrs 11/11; Start tests when cold |
| TBF1 | 1021 | Cold tests |  |
| TBF2 | 1039 | Warm | Was warm 10/11 1800hr 3022 next ? |

In pipeline Next 3034

Test Coordinator this week:  nikolay;  MTM  Tech. on call:  (16-3602) :

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

  - New Instructions: Joints, RRR, Q table u files filling, GHH Sequences /times

  - new todo list  numbers…

  - AOB

SnapShot of STATUS NOW(Tues 11/11am):

SM18 LHC Magnet Test Facility

| TBA1 | _ | _ |  |
| --- | --- | --- | --- |

| TBA2 | MBAR3055 | PT12 Warm Up | Warming |
| --- | --- | --- | --- |

| TBB1 | _ | _ |  |
| --- | --- | --- | --- |

| TBB2 | _ | _ |  |
| --- | --- | --- | --- |

| TBC1 | MBBL1048 | PT12 Warm Up | 300 K |
| --- | --- | --- | --- |

| TBC2 | SSS004 | Prep2 HV Test Warm | 300 K |
| --- | --- | --- | --- |

| TBD1 | _ | _ |  |
| --- | --- | --- | --- |

| TBD2 | _ | _ |  |
| --- | --- | --- | --- |

| TBE1 | MBBR3012 | Prep3 HV Initial @ Warm | 300 K |
| --- | --- | --- | --- |

| TBE2 | MBBR1034 | Prep2.3 IAP @ warm LF | 300 K |
| --- | --- | --- | --- |

| TBF1 | MBBR1021 | MM2 LHC Cycle | 1.9 K |
| --- | --- | --- | --- |

| TBF2 | MBBR1039 | PT13.1 IAP @ Warm | 300 K |
| --- | --- | --- | --- |

Current (pm) Bench Status:

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 04_N | TBE2: Quench in Vclamp Problem already appeared in the morning shift. | MTM |  | Cryo says that everything is OK Exchange of V clamp card with that of TBF2 Origin: Vtap wire disordered in the CFU external to connector. |
| Wednesday 05_M | TBC1: TT821 : no reading because Tsensor coefficients not present in MTF | ACR | MTF |  |
| Thursday 06_A | TBC1: Tcoil : switch in the mobile rack is not changing from MM to Tcoil D1_S09_C opened. | MTM |  | PLC not properly configured? |
| Thursday 06_N | TBC1: Provoked quench all LF : Problem in configuration file | MTM |  |  |
| Friday 07_M | TBF2: Temp probe broken | ACR |  |  |
| Friday 07_A | TBF2: MM0 problem with end stops | MTM |  | Solved by Ewald. |
| Friday 07_A | TBF2: No test type for MM position search included in the test sequence. | MTM | TEMA | Solved by Maryline. |
| Friday 07_A | TBF2: Quench at 1.5 kA (all LF) Extension of quench not properly set by TEMA | MTM |  | Quench repeated |
| Friday 07_A | TBF2: Quench all HF Cryo OK lost while ramping | ACR |  |  |
| Friday 07_A | TBA2: IAP at cold Tcoils. One channel shorted anther shows 431 K | MTM |  |  |
| Friday 07_A | TBC1: training #4 SPA ? Interlock glowing. No trigger, no quench. | MTM/ACR |  | Cryo OK lost ? |
| Saturday 08_M | TBC1: Full load line /ext. joint measurements Message”one bad measurement removed” | MTM | Joint measurement. | Bug in the software Solved by Olav. |
| Saturday 08_M | TBC1: HV at cold Ext&int. Bus Bar v/s ground failed. Test with rack 3. | MTM |  | Re-launched with another rack (rack 2) and OK |
| Saturday 08_A | TBF2: Min energy not performed because Cryo_X2 signal not OK. | ACR |  |  |
| Saturday 08_A | TBF2: Min energy at 11850 A: Configuration file not found.. LF& HF crates switch on/off : MXI error. | MTM |  | LF& HF crates reset. Then OK. |
| Saturday 08_N | TBE2: QH always on when the transparent panel is removed.. | MTM |  | QH voltage not below 50 V. Reset and discharge them. |
| Saturday 08_N | TBF2: Cryo-lost during detraining check. | ACR |  |  |
| Sunday 09_M | TBF2: LHC cycle Program stopped after the cleansing cycle. “Motor end stop and encoder error”. CPU error when receiving the pulse. | MTM |  | Switch off the power supply of the integrators for 10 mms. Problem remaining |
| Sunday 09_M | TBF2: LHC cycle During a ramp rate current measured not readable. | MTM | MMP | Done on purpose during the ramps of current. |
| Sunday 09_M | TBF1: Settings MMP to be done | MTM | MMP | TBF1: Settings MMP updated by Maryline. |
| Sunday 09_M | TBF2: VME error while testing integrators. Integrator cards of the 3rd crate are faulty. | MTM |  | Check by Ewald. |
| Sunday 09_M | TBC1: IAP @warm final , error found in TT821. | MTM |  | No temperature sensor in the magnet of TBC1? |
| Sunday 09_M | TBF2: Final HV insulation @warm : change of the switch (CC to F2) Position.. Switch blocked. | MTM |  | Maryline carried out the change. |
| Sunday 09_N | TBA2: After MM: MMP is hung | AB/CO | MMP | Reboot the sun to continue |
| Mon 10_M | TBF1: MM position search MUX , VME, motor errors. | MTM |  | Test aborted. |

Problems of Last week

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 04_N | TBE2: Quench in Vclamp Problem already appeared in the morning shift. | MTM |  | Cryo says that everything is OK Exchange of V clamp card with that of TBF2 Origin: Vtap wire disordered in the CFU external to connector. |
| Wednesday 05_M | TBC1: TT821 : no reading because Tsensor coefficients not present in MTF | ACR | MTF |  |
| Thursday 06_A | TBC1: Tcoil : switch in the mobile rack is not changing from MM to Tcoil D1_S09_C opened. | MTM |  | PLC not properly configured? |
| Thursday 06_N | TBC1: Provoked quench all LF : Problem in configuration file | MTM |  |  |
| Friday 07_M | TBF2: Temp probe broken | ACR |  |  |
| Friday 07_A | TBF2: MM0 problem with end stops | MTM |  | Solved by Ewald. |
| Friday 07_A | TBF2: No test type for MM position search under the test sequence. | MTM | TEMA | Solved by Maryline. |
| Friday 07_A | TBF2: Quench at 1.5 kA (all LF) Extension of quench not properly set by TEMA | MTM |  | Quench repeated |
| Friday 07_A | TBF2: Quench all HF Cryo OK lost while ramping | ACR |  |  |
| Friday 07_A | TBA2: IAP at cold Tcoils. One channel shorted anther shows 431 K | MTM |  |  |
| Friday 07_A | TBC1: training #4 SPA ? Interlock glowing. No trigger, no quench. | MTM/ACR |  | Cryo OK lost ? |
| Saturday 08_M | TBC1: Full load line /ext. joint measurements Message”one bad measurement removed” | MTM | Joint measurement. | Bug in the software Solved by Olav. |
| Saturday 08_M | TBC1: HV at cold Ext&int. Bus Bar v/s ground failed. Test with rack 3. | MTM |  | Re-launched with another rack (rack 2) and OK |
| Saturday 08_A | TBF2: Min energy not performed because Cryo_X2 signal not OK. | ACR |  |  |
| Saturday 08_A | TBF2: Min energy at 11850 A: Configuration file not found.. LF& HF crates switch on/off : MXI error. | MTM |  | LF& HF crates reset. Then OK. |
| Saturday 08_N | TBE2: QH always on when the transparent panel is removed.. | MTM |  | QH voltage not below 50 V. Reset and discharge them. |
| Saturday 08_N | TBF2: Cryo-lost during detraining check. | ACR |  |  |
| Sunday 09_M | TBF2: LHC cycle Program stopped after the cleansing cycle. “Motor end stop and encoder error”. CPU error when receiving the pulse. | MTM |  | Switch off the power supply of the integrators for 10 mms. Problem remaining |
| Sunday 09_M | TBF2: LHC cycle During a ramp rate current measured not readable. | MTM | MMP | Done on purpose during the ramps of current. |
| Sunday 09_M | TBF1: Settings MMP to be done | MTM | MMP | TBF1: Settings MMP updated by Maryline. |
| Sunday 09_M | TBF2: VME error while testing integrators. Integrator cards of the 3rd crate are faulty | MTM |  | Check by Ewald. |
| Sunday 09_M | TBC1: IAP @warm final , error found in TT821. | MTM |  | No temperature sensor in the magnet of TBC1? |
| Sunday 09_M | TBF2: Final HV insulation @warm : change of the switch (CC to F2) Position.. Switch blocked. | MTM |  | Maryline carried out the change. |
| Sunday 09_N | TBA2: After MM: MMP is hung | MTM | MMP | Reboot the sun to continue |
| Mon 10_M | TBF1: MM position search MUX , VME, motor errors. | MTM |  | Test aborted. |