vc/20-Feb-04

### Week 06  starting 02 Feb 04

### SM18 Operation Meeting (Tuesdays)

Schedule for now  + see Display :

|  | Forward look, & Next Magnets |
| --- | --- |
| TBC1 MBBR2041 standby TBC2 SSS08 Cold tests 1.9k TBD1 MBAL1063 Cold tests 1.9 K TBD2 MBBL3074 disconnect warming TBE1 CFB repair NEW MAGNET WEDNESDAY ? 300 K TBE2 MBBL3076 1068 next with Correctors to Connect TBF1 MBAR1053 Pump down +cooldown TBF2 MBBR2044 ARRIVED Mon. 2Feb 300 K | 3082 2073 1033 3079 1044 |

Test Coordinator this week: V.Chohan + John Miles;   MTM  Tech. on call:  (16-3602) :

Announcements

- Rules Flowchart

- ICS To be CALLED UNDER ALL CIRCUMSTANCES WHEN REQUIRED

Tel. Numbers in SMCR Grey folder + Board

Shift Leaders know the latest situation/prediction cryo etc & WILL do so (despite French etc)

- SIGNING Traveller at the end :  Remove 2A, & already sign for REMOVAL even if shafts in	Because Cryo Will not break vacuum till shafts are out ( problem Sun night/Mon)

- Ohmic manual M’ments

Templates in New Home page ( Mahesh) ; Link there in old

Start saving in same directory as MTF HV files ( in dfs  ) for ease of upload to MTF

- New : CDPT file to be EXCEL  created from SMTMS  ( will include minimal Quench Table	Automatically as well as MAPS in future .. ( in progress);  	This FILE must be UPLOADED to MTF

PROBLEMS OF LAST WEEK

START UP PROBLEMS are Normal but……

SUPPORT & SERVICES

- More than 36 hrs LOST due to LEAK ( CFB-to short antic –MTM Mech) : Lack of Organized Service in AT-MTM

- Power Supply Repair ( ~ 48 hrs lost) : D Cluster : We need to Organize AB-PO Service

- Shafts Service in MTM needs to be organized – Problem is not in Shafts Insertion but in taking Out ( which is needed day or night, Sat or Sun depending on tests finishing !) ; ICS & CRYO interaction Hangs up

- Rotary Switch Issues & MTM Eq Support to do it loses several hours all the time ( hopefully fixed this week)

- Hardware Recognition System Problems causes certain delays…. Consolidation Required etc

# Main Events during last week.

Cryogenic  and vacuum (during connection and test):

- Cold compressor restart only Thursday evening (only 6g/s before)

- Oil leak in one pump (stop Wednesday).

- Problems in the pumping system (automate, Temperature too high..)

- leak in the M2 line joint for SSS08 during installation.

- leak in the big below joint for 2041 after the closing of the magnet.

- leak in the short anticryostat for F1 (ap2) : MTM intervention needed.

Power :

- Switch (cluster F, no motorisation, inversion of warnings)

- Resistance that burns in the Power Converter of cluster D after a quench:

Control :

- Chassis VME not working in cluster D

- HRWC not working properly during the night ( problem with the DB, to discuss with TS/CSE) .

- Applications cluster C that had to be re-implemented.

- Problem in RS 232 modules : not the same for all TRU.

Problems  of “ MTM” Type: see the detailed list next slide.

DETAILED FAULTS FOLLOW-UP ( Test Coord  + Stephane S)

Problems of Last week

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Monday 26_M | TBF1 : IAP at cold TBF1: No access in Vol 7 (backup of the files in the control room) | MTM | IAP | Done in local. |
| Monday 26_N | TBC2: SSS08 PLC not working No screen of the sun and the applications present to perform electrical test at warm | AB/CO AB/PO | All | PLC repaired by C.Dehavay Sun screen back Monday afternoon. HF and LF introduced by H.Reymond Tuesday morning. |
| Monday 26_M | TBD1: IAP@ warm (1063) EE014 and EE015 not connected. | MTM |  | Magnet coming without diode. EE013 used as reference point. |
| Monday 26_A | TBF1: Loss of cryo OK Leak of oil in warm roots. Temperature in the line is increasing. | ACR |  | Stop Wednesday morning. Restarted after 7h of stop |
| Monday 26_A | TBD1: IAP @warm: Interlock problem | MTM | IAP | F.Chiusano contacted. Solved |
| Monday 26_A | TBD2: IAP @cold: HF@LF acquisition could not be armed. | MTM |  | Application closed and started again. |
| Monday 26_A | TBD1: IAP @warm: all LF No led is glowing in the inter-connexion box. No bench selected | MTM | IAP |  |
| Monday 26_A | TBF1: IAP @ cold (Tcoil): Missing equipment. TRU not scanned. | MTM | HWRC | Scan performed but the problem persists, |
| Monday 26_N | TBF1: Only shaft in ap2 but HWRC expected 2 apertures | MTM | HWRC | Scan stopped. Software modified by M.Gateau OK now. |
| Tuesday 27_M | TBF1: Rotary switch problem in position 2. Warning appeared in the other bench. | AB/PO |  | G. D’angelo organised the modification. |
| Tuesday 27_M | TBD2: IAP at warm done in local. No access in Vol 7 | MTM |  |  |
| Tuesday 27_M | TBF1: No access in Vol 7 for LFOFF | MTM |  | OK now |
| Tuesday 27_M | TBF1: Problem with connector of TRU2 (DB25 mal for RS232 not fit with connector of TRU DB9) MM1 has to be done after the training | MTM |  | TRU 2 not ready. In the future operational TRU will be marked. |
| Tuesday 27_A | TBD2: IAP at cold : Wrong magnet name appeared. | MTM |  | TEMA restarted by M.Gateau. |
| Tuesday 27_A | TBF1: MM1 Error motor in auto test | MTM |  | Shaft in one aperture: auto test performed only in aperture 2. |
| Tuesday 27_A Tuesday 27_N | TBD2: IAP at cold : No data when HF is triggered. | AB/CO | IAP | VME power supply, chassis changed Wednesday night. |
| Tuesday 27_A Tuesday 27_N | TBF1: LHC cycle :Error missing equipment. During the tests VME error appeared and no data was present. No trigger was sent by the encoder to the integrator. Same problem for the load line. Config. file nit changed for measurements in one aperture. | MTM |  | LHC cycle cancelled. Load line done in manual mode downstairs with Federico. |
| Tuesday 27_N | TBF1: Min energy 11850 for warm up TEMA not procceding. HWRC shows error while getting the equipment in the database. | MTM |  | HWRC disabled. |
| Tuesday 27_N | TBF1: Min energy 11850 for warm up Cryo loss.: slow PA abort | ACR |  |  |
| Tuesday 27_N | TBD2: HV test : problem in MOBIC: problem in Microsoft software. |  |  | F.Chiusano contacted the helpdesk Data has to be entered manually because the Laser Reader is not working. |
| Wednesday 28_M | SSS08 not closed because a leak is suspected in the M2 line. | ICS |  | Joint changed by ICS. |
| Wednesday 28_A | TBE2 (3076) : IAP @cold Offset cannot be adjusted better than 122 mv for D1-L-U | MTM | IAP |  |
| Wednesday 28_N | TBE2 (3076) :HV test @cold File could not be uploaded. CDBP file not created. | MTM | HV | Repeated and then OK |
| Wednesday 28_N | TBE2 (3076) : Cannot launch slow PA Gammelle not scanned | MTM | HRWC |  |
| Thursday 29_M | TBE and TBF: HRWC problem with TEMA. No magnet name. | MTM | HRWC | Solved by M. Gateau |
| Thursday 29_M | TBD2: IAP @cold. Measurements through the patch panel not stable. Difficult to adjust the offsets. | MTM | IAP |  |
| Thursday 29_M | TBC1: Warm tests in C1. Protection in quadrupole configuration LF charged.. Changes in the procedure for the re-configuration of the protection for the dipole. | MTM |  | Message to be sent to all equipment supports. Procedure to be written. |
| Thursday 29_A | TBC1: IAP @cold. Files created with the wrong extension (.x1 instead of .c1) | AB/CO | HRWC | Corrected by A.Raimondo |
| Thursday 29_A | TBC1: Qloc cards not cabled | MTM |  |  |
| Thursday 29_A | TBD: No communication for powering Unable to launch hfoff from sun MTA35 Quench at 1.5 kA | AB/CO | hfoff | hfoff reinstalled by P.Coutinho |
| Thursday 29_A | Sun MTA37 : access denied. Unable to get MMB pens | AB/CO |  | M.Bjork contacted. Solved. |
| Thursday 29_A | TBD2: Quench provoked at 1.5 kA Faulty. Symmetric quench Vbus triggered before | MTM |  | OK given by test coordinator. |
| Thursday 29_A | TBD1: Unable to launch cool down due to anticryostat problem Wrong cabling. | MTM |  | Solved by Technical support. |
| Thursday 29_N | TBE2: MM1 TRU problems, (powering inversed) MUX error in card1, Motor serial comm. Error because of RS communication modules non standard Shaft alignment error (sectors in overranges) Read current zero | MTM/AB-CO |  | RS communication modules non standard depending on the TRU. To be standardized by (AB/CO). . |
| Thursday 29_N | TBE2: HRWC error while launching the training quench | MTM | HRWC | Solved by M.Gateau Friday morning |
| Friday 30_M | TBE2: Training quench number 2 No file for Qloc. No switch from MM to Q.loc? | MTM |  |  |
| Friday 30_M | TBD2:Incomplete setting for TEMA for cluster F in vol7 | MTM |  |  |
| Friday 30_A | All benches : Lost of the cryo OK. Problem in one automate related to the pumping system | ACR |  | Restarted after 1h of stop |
| Friday 30_A | TBE2: MM1 no current reading | MTM |  | Integrator racks left in PDI configuration. |
| Friday 30_A | TBE2: Leak in the big below for 2014. | ICS |  | Reopened by ICS and closed again. |
| Friday 30_N | TBF2 and TBD: Power Converter : Main PC fault | AB/PO |  | Solved for TBF. TBF2: Interlock due to the manual connection used when the motorised switch cannot be used. TBD: Solved Sunday night by H.Thiesen. Resistance of power converter circuit protection out of work. |
| Friday 30_N | TBF2 : Provoked quench all LF. Messages trigger received and file name did not appear in the HF acquisition window | MTM | Power | Power application window closed and messages appeared. |
| Friday 30_N | TBF2 : Provoked quench all HF .HRWC in TEMA gave error in DB Id for the training quench. | MTM/TS-CSE? | HRWC | HRWC recognition by-passed. |
| Friday 30_N | TBC: anticryostat not OK. Cool down could not start. Wrong parameters | MTM |  | Parameters changed by P.Legrand Saturday evening. |
| Saturday 31_M | TBE2: Polarity interlock when launching power application HRWC did not work in TEMA. | MTM/TS-CSE? | HRWC | HRWC disabled. |
| Saturday 31_M | All benches : Lost of the cryo OK. Problem in the pumping group. | ACR |  | Restarted after 5 h of stop |
| Saturday 31_M | TBC1 : Problem of communication with the PLCSE. User name not valid. | AB/CO |  | Problem solved Monday morning by C.Dehavay. |
| Saturday 31_M | TBE2: MM1 problem with end sectors of ap 2. | MTM |  | Switch to e-coil. No results. Change of the mobile rack n2 instead of n3. The MUX problem. DC power supply transferred from rack 3 to rack 2 because channels were missing in power supply of rack n2. Then OK. Rack n3 has to be controlled before another using Rs232 com has to be checked in the cluster E. |
| Saturday 31_M | TBE2: Training quench at 4: No Qloc. Switch from MM to Qloc forgotten. | MTM |  |  |
| Saturday 31_N Sunday 01_M | All benches : No cryo OK. Problem in the pumping group. | ACR |  | Restarted after 5 h of stop |
| Sunday 01_M | TBE2: LHC cycle : HRWC not working. TEMA blocked | MTM/TS-CSE? | HRWC | HRWC disabled. TEMA restarted. |
| Sunday 01_N | TBF1: Leak in the short anticryostat (Ap2). Cool down blocked by test coordinator. | MTM |  | Problem fixed Monday morning because no mechanical technicians available. |
| Monday 02_M | TBD2: HEWC problem during the launch of training quench | MTM/TS-CSE? | HRWC | Solved by M.Gateau. |