vc/20-Feb-04

### Week 07  starting 09 Feb 04

### SM18 Operation Meeting (Tuesdays)

Schedule for now  + see Display :

| Bench | Name | Test type | Temp | - | - |
| --- | --- | --- | --- | --- | --- |
| TBC1 | MBBR1044 | START TESTS WED NIGHT? |  |  |  |
| TBC2 | SSS008 | FINISH WED Midday ? | 1.9 K |  |  |
| TBD1 | MBBR1037 | START WHEN D2 finished? |  |  |  |
| TBD2 | MBAR2013 | Sunday's problem? | 1.9 K |  |  |
| TBE1 | MBBR1030 | NEW One in today ? |  |  |  |
| TBE2 | MBAR3030 | ICS1 WP04 HV Test Warm |  |  |  |
| TBF1 | MBBL3082 | START WHEN F2 finishes |  |  |  |
| TBF2 | MBBR2044 | PT 5.2 1.5 KA Quench HF | 1.9 K |  |  |

Test Coordinator this week: Stephane ( + Gijs de Rijk);   MTM  Tech. on call:  (16-3602) :

General Points

- CRYO Issues ( Jean-Paul)

- ICS Leaks ( a lot over weekend: B Poulnais knows; Glue problem)

- Problem on 2013 on Sun night ( Guy)

- SSS & Switch over ( Guy, Nikolay) for Wed .

- D1 work on CFB ( Lionel) for corrector leads ?

- PLC CU tests etc for Claude Dehavay ?

- Do not forget PT8  [ramp to 12000A & re-ramp at 120A/sec etc]

- The Must List for Magnetic Measurements ( only 4 remain now)

- ICS To be CALLED UNDER ALL CIRCUMSTANCES WHEN REQUIRED ; they work 24 hrs/day , 7 days/week	Tel. Numbers in SMCR Grey folder + Board	Shift Leaders know the latest situation/prediction cryo etc & WILL do so (despite French etc)

- SIGNING Traveller at the end :  Remove 2A, & already sign for Standby ( if non-conformitites)  or stripping even if shafts in 	Because Cryo Will not break vacuum till shafts are out ( problem Sun night/Mon)

- Ohmic manual M’ments

New Templates in New Home page ( Mahesh) ; Link there in old

Start saving in same directory as MTF HV files ( in dfs  ) for ease of upload to MTF

- New : CDPT file to be EXCEL  created from SMTMS  ( will include minimal Quench Table	. ( in progress);  	This FILE must be UPLOADED to MTF

PROBLEMS OF LAST WEEK

DETAILED FAULTS FOLLOW-UP ( Test Coord  + Stephane S)

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday_03_M | TBD1: TRU 3 problem due to RS232 communication | AB/CO |  | Solved during the morning |
| Tuesday_03_A | TBC1: HV test login refused on sunmta19. Password file missing | AB/CO | HV | Solved |
| Tuesday_03_A | TBD1: Cabling not correct in the mobile rack of ap1 coil signals | AB/CO | HV | Solved |
| Tuesday_03_N | TBD2: TT821 out of order | MTM |  |  |
| Tuesday_03_N | TBC1: Water: valve EV did not follow the command of the PLCSU | MTM |  | Opened manually. Solved by Federico |
| Tuesday_03_N | TBD2: Helium not flowing in the anticryostat | MTM |  |  |
| Tuesday_03_N | TBD1:TRU, mobile rack not scanned at proper time. Qloc not recorded during the training quenches, | MTM |  |  |
| Tuesday_03_N | TBD1:TRU, mobile rack not available in the database. LHC cycle stopped by HWRC | MTM | HRWC |  |
| Tuesday_03_N | TBD1:Re-scanning of the missing data not possible because of a network problem. Site not accessible by the mobic | MTM | HRWC |  |
| Tuesday_03_N | TBC1: connector D20 not plugged back after HV Trigger sent when we start to power magnet. | MTM |  |  |
| Tuesday_03_N | TBC1: SPA: HRWC system gave error: while getting equipment list from MTF in the database.. | MTM | HRWC | HRWC disabled and SPA re-launched. |
| Wednesday_04_A | TBD1: last quench before warm up ; problem with LF trigger and I-earth signal | MTM | LF | Retried. Then OK. |
| Wednesday_04_A | EVC cannot be controlled automatically. The valve is broken and is under repair | MTM |  | Handle in manual mode.. |
| Wednesday_04_N | TBF1: During a quench AQA give error of 1% in tau on one heater. | MTM | AQA | Maryline contacted. LF signals checked. OK. Acquisition error. |
| Wednesday_04_N | TBF1: QH YT221 had a lower resistance than the errors | MTM |  | HV test done then OK. |
| Tuesday_05_M | TBD2: Problem when QH charged because permanent trigger present. Cable 20 V, 18 A not properly connected (socket missing). | MTM |  | Maryline called and repaired. |
| Tuesday _05_A | TBD2: Intercon box not working. No led glowing in the panel. CTH sent information to PLCSE that magnet was below 20 K and it was warm. | MTM | ACR/AB-CO? | A.Tovar called for repair |
| Tuesday _05_A | Mobile rack in wrong state when receiving order from TEMA. Switch in FA position and not in Qloc. | MTM |  |  |
| Tuesday _05_N | TBE2: IAP at cold trigger on potaim cards Vbus ext-int. | MTM |  | Solved by Maryline |
| Friday _06_M | TBF2: Control of the rotary swithch : no led was glowing | MTM |  | Maryline contacted |
| Friday _06_M | TBF2: MMP launched : problem of amplifier. | MTM | MMP | Aborted. |
| Friday _06_N | TBC1: HRWC scan for HV XMT.MY92 does not exist. | MTM | HRWC |  |
| Friday _06_N | TBE1: Leak discovered during leak tests.. | ICS |  | ICS called. |
| Saturday _07_M | TBC1: Leak on anticryostat | ? |  |  |
| Saturday _07_M | TBD1: Leak on the below | ICS |  |  |
| Saturday _07_A | TBE2: IAP final at warm: alarm on anticryostat X2 | MTM |  | Guy reset the PLC |
| Saturday _07_A | TBE2: CDW2 non conformity Corrector resistance too high | MTM |  |  |
| Sunday _08_A | TBE2 : Slow PA during ramp on the fifth training quench. At a current of 5171 A The HV test performed between dipole and ground and it failed. | MTM |  | Cryo OK. Problem coming from the magnet? Investigation done Monday. Everything OK. |
| Sunday _08_A | TBE1:Slow PA could not be launched. Cryo problem. | ACR |  |  |
| Monday_09_M | TBD2: Permanent trigger. Files from a05 to a049 to be erased. | MTM |  | Acquisition is kept off |
| Monday_09_M | No PT8 performed for 1068 and 1030, 1053, 2044.. PT8 mandatory in each magnet. | MTM |  |  |
| Monday_09_M | TBE: cables running from the mobile rack damaged by the ROCLA? | ICS? |  | Repaired by P.Legrand |