Vc/14-Apr-04

### Week 16  starting 12April 04

### SM18 Operation Meeting (Tuesdays)

Bench status  and tests for now & next few days :

| Bench | Name | STATUS Mon 12/4 Afternoon | Temp | - | - |
| --- | --- | --- | --- | --- | --- |
| TBB1 | MBAL1059 | PT 5.2 1.5 KA Quench HF | 1.9 K |  |  |
| TBB2 | MBBL3119 | PT 12 Warm Up | Warming |  |  |
| TBC1 | ? | New Tues 13/4 |  |  |  |
| TBC2 | SSS0013 | EX1 Extra Test | 300 K |  |  |
| TBD1 | MBAR3053 | ICS 3 Traveller after warm test | 300 K |  |  |
| TBD2 | MBAL1073 | PT 12 Warm Up | Warming |  |  |
| TBE1 | 3088? | New Tues 13/4 |  |  |  |
| TBE2 | MBBL3117 | ICS 3 Traveller after warm test | 300 K |  |  |
| TBF1 | MBAR2019 | Prep 5 Pumping + Cool Down | Cooling |  |  |
| TBF2 | 3057 ? | NEW Monday 12/4 |  |  |  |

Test Coordinator this week: Gijs de Rijk ; MTM support on call  Richard:  (16-3602).

- Some statistics

11 Dipoles  Disconnected/removed  after tests last week  ( Sat Midnight to Sat midnight)  RECORD

[ Previous Week was 5 Dipoles ]

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- General Problems last week

| sm18 cooling water | not enough for > 3 converters, delays experienced on 3118(D1) on Monday |
| --- | --- |
| Vacuum leaks | big delays on F1-2019 (two leaks, 1.5 days lost); two more leaks repaired on C1 tue morning (He lines) and sat afternoon (cryostat) |
| ICS overtime | Tue eve: as ICS personnel already run out of their overtime quota for the week, 3058 could not be connected on E1 and 1062 could not be disconnected from D2 |
| HV rack for ICS | software problem, Ewald and Adriaan unsuccessfully attempted fix on Thu afternoon. During all weekend, MTM operators carried out warm HV test using own rack, which lacks connectors for aux busbars -> aux busbars of 1059, 2019, 3117 and 3053 must be re-tested in SMA18 as soon as ICS rack is repaired |
| Cryogenics | Regeneration of heat exchanger done wed afternoon as agreed, 8 hors lost for cooldown/warmup in gas phase; In addition, 6 hour lost when warming up 1057 due to an electrical short circuit found and repaired during the regeneration |
| Cryogenics | 12 hours lost during the cooldown of 3111 on Friday morn, due to unspecified problem |
| MTM-ICS coordination | Saturday: delay for the connection of 3117 on E2, ICS were called only late evening |
| Old SSS06 removed from C2 | Parked on the 3 temporary feet in font of B1 bench, impossible to remove B1 before Tuesday (sm18 transport operators needed to operate crane) The feet, even w/o magnet, make it very difficult the transport of dipoles from/to B1 (problems experienced on wed aft) |
| SSS13 | Blocked on C2 all weekend CONNECTION NOT DONE : needs MTM Mech work/alignment before ICS can continue |

Other General Issues

  - D Cluster Commissioning of 600A & 150 A Power Supplies for future SSS setup on this cluster ( Guy + Giorgio) : next week?

  - Keep D1 free for SSS Mech Preparation ( 6 hrs ? )   J.Halik

[From 19 April…]

| Date | Bench | Problem | Type | Application/soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Wednesday7_A | All | Regeneration of heat exchanger. It took 8 hours. In addition 6 hours lost due to an electrical short circuit. | ACR |  | Repaired during regeneration then OK |
| Thursday 8_M Thursday 8_A Friday 9_M | TBF1 | Software problem with the Megger of ICS during HV test. During HV test, error in putting in REMOTE.. For the magnet 1057 WP7_HV not available. | AB/CO? | HV | HV Test at warm (WP04, WP07) done with MTM rack. But test of bus bar not performed (no cables). HV test of bus-bars must be performed for 1059,2019,3117 and 3053. Quit the program and restart then OK. Problem of transferring the file? |
| Friday 9_M | TBF1 TBE2 | 2019: New magnet. HV test at warm (WP04). Test failed. Same problem for HV test (WP07) | MTM MTM | HV | Ewald called. He re-started the test several times and then OK. Ewald called. Tests perform in the D2 sun station connected to E2 with the dedicated pair of cables. Warning : Report generated in D2! |
| Friday 9_M | TBF2 | Cool-down of the 3111 too long (12 hours lost) . Origin not found. | ACR |  |  |
| Saturday_10_M | TBF2 | After a quench, difficulty to recover 1.9 K. Magnet blocked at 5 K. | ACR |  | ACR piquet called. Level temperature probe problem? |
| Saturday_10_N | TBC1 | SPA at 1 kA. Power supply error. No ramp. The status of the PC was “locked” and from sun MTA24 it was not possible to launch the server. | MTM |  | Ewald called. PC was in SSS configuration. Ewald put the PC in correct configuration for a dipole. |
| Saturday 10_N | TBD1 | HV test (WP04) could not be launched. | AB-CO? |  | It worked after re-booting the sun. |
| Sunday_11_A | TBD1 | HV at warm (WP04): file not uploaded because not located. | MTM |  |  |
| Sunday_11_A | TBB | Access to the rotary switch blocked by a pump | ACR |  |  |
| Sunday_11_A | TBE1 | IAP @warm final : test launched from TEMA. Correct bench selected but interconnection box in TBE2. 2.5 A sent to TBE2 by mistake. File generated in folder of TBE1. | MTM | TEMA | Test repeated re-selecting the bench on local. The results of the two files were identical. |
| Sunday_11_A | TBD2 | Magnet 1073 reached ultimate field. During the 10-minutes plateau the current went down and went up to 12850 A. Then it stayed at 12850 A. | MTM |  |  |