Vc/27-May-04

### Week22 starting 24 May 04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Sunday 23/5 due to ROCLA problems:

| Bench | Name | Status Sun 23May04 21.30hr | Temp |
| --- | --- | --- | --- |
| TBA1 | MBAR3061 | Prep 5 Pumping + Cool Down | Cooling |
| TBA2 | MBBL3117 | PTE 27.1 DRQ 4K | 4.4 K |
| TBB1 | MBBL1104 | PT 6.2 Training 2 | 1.9 K |
| TBB2 | MBAL1071 | waiting for Rocla | 300 K |
| TBC1 | MBBL1116 | PT 11 4 K Quench SSL | 4.4 K |
| TBC2 | SSS21 | PT 6.3 Training 3 | 1.9 K |
| TBD1 | MBBL2076 | waiting for Rocla | 300 K |
| TBD2 | MBBL2062 | PT 6.4 Training 4 | 1.9 K |
| TBE1 | MBBL1112 | waiting for Rocla | 300 K |
| TBE2 | MBBL1113 | waiting for Rocla | 300 K |
| TBF1 | MBBL3151 | ICS 5 Traveller Last | 300 K |
| TBF2 | MBAL3160 | waiting for Rocla |  |

LOTS FOR PBLMS : PLC, HWREC etc etc, etc last long holiday weekend

ROCLA were out of USAGE from Wed 19th May onwards thro’ long weekend

Bench status  and tests for now ( Monday 24/5 aft)  & next few days :

| Magnets under Test on 24-05 4:10:01 PM |
| --- |
| - | Bench | Magnet | Shafts | Start | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | Last Test Desc | Time | I | Temp. | Total |
| 1 | TBA1 | MBAR3061 |  | 14-05 | 146 h. | 108 h. | 2 h. | - | - | PT 2 HV Cold Test | 24-05 (14 : 10) | 0 A. | 1.9 K | 256 hours |
| 2 | TBA2 | MBBL3117 |  | 12-05 | 112 h. | 25 h. | 149 h. | - | - | MM 2 LHC Cycle | 24-05 (14 : 12) | 0 A. | 1.9 K | 286 hours |
| 3 | TBB1 | MBBL1104 |  | 17-05 | 26 h. | 35 h. | 111 h. | - | - | PTE Min. NRJ 3KA Diode Test | 24-05 (15 : 24) | 3000 A. | 1.9 K | 172 hours |
| 4 | TBB2 | MBAL1071 |  | 11-05 | 52 h. | 55 h. | 23 h. | 13 h. | 21 h. | waiting for Rocla | 18-05 (15 : 33) | 0 A. | 300 K | 309 hours |
| 5 | TBC1 | MBBL1116 |  | 17-05 | 59 h. | 42 h. | 49 h. | 19 h. | - | PT 12 Warm Up | 23-05 (21 : 50) | 2 A. | Warming | 169 hours |
| 6 | TBC2 | SSS21 |  | 13-05 | 14 h. | 82 h. | 174 h. | - | - | PT 6.3 Training 3 | 18-05 (01 : 00) | 12458 A. | 1.9 K | 270 hours |
| 7 | TBD1 | MBBL2076 |  | 13-05 | 14 h. | 58 h. | 42 h. | 17 h. | 31 h. | waiting for Rocla | 20-05 (02 : 17) | 0 A. | 300 K | 272 hours |
| 8 | TBD2 | MBBL2062 |  | 17-05 | 46 h. | 66 h. | 53 h. | 5 h. | - | PT 12 Warm Up | 24-05 (11 : 30) | 0 A. | Warming | 170 hours |
| 9 | TBE1 | MBBL1112 |  | 12-05 | 142 h. | 0 h. | 21 h. | 13 h. | 4 h. | waiting for Rocla | 19-05 (20 : 00) | 0 A. | 300 K | 296 hours |
| 10 | TBE2 | MBBL1113 |  | 11-05 | 51 h. | 38 h. | 58 h. | 16 h. | 35 h. | waiting for Rocla | 20-05 (02 : 18) | 0 A. | 300 K | 308 hours |
| 11 | TBF1 | MBBL3151 |  | 17-05 | 49 h. | 31 h. | 38 h. | 16 h. | 38 h. | Finished Ready to go Final scan | 24-05 (15 : 29) | 0 A. | 300 K | 173 hours |
| 12 | TBF2 | MBAL3160 |  | 11-05 | 46 h. | 79 h. | 46 h. | 15 h. | 11 h. | waiting for Rocla | 19-05 (15 : 00) | 0 A. | 300 K | 318 hours |

Test Coordinator this week: Gijs de Rijk ; MTM support on call :  (16-3602).

4 Dipoles  Disconnected after tests last week    ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

A2  commissioning   & re-test     3117

A1         “                   &   re-test    3061

- D1 &  preparation mechanically for SSS

- SHADOW ACTIVITIES DUE TO ROCLA Problems

  - Install Rotary Switch in E Cluster ( 5 days work asked)

  - Repair Water leak in F cluster Power Coverter ( 3 days asked)

  - Calib B Cluster Power Supply in cold condition  : needs to be done once

  - Calib A     “           “            “      “     “        “               “        “          “       “

- SSS work :        To Do list on Board+ GREY FILE IN SMCR

General Problems last week :

| Date | Bench | Problem | Type | Application/soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday 18_N | TBA2 | MM: No possibility to power the magnet either from TEMA or locally. | AB-CO? | TEMA, MMP, Power |  |
| Wednesday _19_N | ALL | Cryo Display is working badly. Communication problem | ACR | Display | Federico checked if the anticryostat coefficients were OK in the PLCart.. |
| Wednesday _19_N | TBA2 | Before starting LHC cycles, HWREC found that the SUN station was not scanned. | MTM |  | Sun scanned with the help of Federico. |
| Wednesday _19_N | TBA2 | LHC cycle not exported. Magnet name format error | MTM | AB-CO? | Federico called. Name of the cryodipole replaced by the name of the cold mass by Federico. He did the same for TBB1. The problem remained for ramp in cluster A. Stephane advised to note down the measurements which could not be exported .He will contact AB-CO Monday. |
| Wednesday _19_N | TBA2 | During magnetic measurements MMP stopped because of communication errors. | AB-CO | MMP | Auto-test performed in MMP and then tests re-started. |
| Thursday _20_M | TBB1 | Quenches at 1.5 kA: HF DAQ could not be armed . Origin : MXI error. | AB-CO | HF | TEMA quitted and re-started: no effect. Another trial one hour after, then OK. |
| Thursday _20_M | TBA2, TBB1 | The 3 new LHC cycles in the extended program implemented in the cluster F were not present in the other clusters. | MTM |  | Stephane called Federico. He copied the three cycles from the cluster F to the cluster A and B. and to the other clusters. But He did update the tests sequences only for A and B. The update has to be performed for the other clusters in the future (Stéphane &Maryline). |
| Thursday _20_A | TBA2 | Ramp test on A2: Power converter sent communication error. | AB-PO |  | Federico checked the interlocks, reset the communication card in RMA and start the PC with Power in local. Then GHH launched Power remotely and then launched the test. |
| Thursday _20_A | TBB2 | Cables inverted between B1 and B2 in the trenches. The 2 cables with the 104 pins which are connected to the TRU are missing? | MTM |  | Federico fixed the problem. |
| Thursday _20_A | TBB1 | Magnetic measurements performed with the Head 24. The head is not defined in MTF. | MTM | MTF | Maryline has to contact EDMS support. Federico entered the heads 27 instead the heads 24 to let the test be performed. |
| Thursday _20_A Friday_21_M | TBB1 | HWREC informed that config. File contained Oracle errors. |  |  | He came to edit the files manually. |
| Thursday _20_N | TBA2 | During LHC cycle FT 60s PC main fault Power supply parameter ‘active’ was off. |  |  | . |
| Thursday _20_N | TBA2 | MM1: TRU motor is not working correctly. It gave communication problems.. | MTM |  | . |
| Thursday _20_N | TBF1 | PTE: no response from the PLCSE. | MTM |  | . |
| Friday_21_M | TBB1 | 2 cables numbered 880 and 882 used to bring 9 V to the RS232 port did not work. | MTM |  | Federico replaced them by the 881 and 883 cables. He will inform Maryline about the change. |
| Friday_21_M | TBF1 | TEMA is not displaying the good status of the PLCSE. | AB-CO | TEMA | Federico restarted TEMA and the server application of the SUN, then OK. |
| Friday_21_M | TBA2 | A worldfip cable was disconnected from the front face of RMA. Launching TEMA MMP gave the following error “the RMB is giving communication error” (why RMB?) | AB-CO | MMP/Power | Federico checked the power application, he checked that the RMA is locked by the right SUN. He phoned to A.Rijllart and launched MMP in local mode. It worked so he suggested to continue the tests in this way until Monday. |
| Friday_21_A | TBC2 | Switch to SSS tests: Access error “access error ,device is locked” when Federico tried to change the magnet parameters in the application Power. | AB-CO? | Power | Power supply selection corrected in power application. |
| Friday_21_A | TBB2 | MM position: no results because mobile rack not switched in MM position. Problem in the rotation of the shaft in aperture 2: TRU in aperture 2 has un untight clutch. | MTM |  | Federico has tight the clutch. |
| Friday_21_A | TBB1 | In matest no magnet folder template valid for MBAL magnet for B bench. Federico upload the files generated by HWREC but without success. | MTM |  | Test stopped. Priority given to the other benches. |
| Friday_21_A | TBA2 | MBAL 3117: Iearth high: 19 mA.. | MTM |  |  |
| Friday_21_N | TBC1 | PC main fault : access error on the PLCSE. | MTM |  |  |
| Saturday_22_M | TBA2 All benches | LHC cycle could not be launched from TEMA as well from local server. Message : “application with the same name running” The message came from local server for power despite all applications were closed. | AB-CO? | TEMA, MMP | Federico managed to run the tests on local SUN. The tests have to be performed on this way until Monday and the intervention of AB-CO to solve the problem Federico asked for a clear statement on the status of the X-console window. For most of the SUN the window is not opened. Why? |
| Saturday_22_M Sunday_23 _M | TBC1 | Provoked quench at 1.5 kA (LF heaters). HF DAQ did not arm. Test repeated but no data in AQA. HF acquisition did not receive a trigger. Same problem occurred during a training quench. | AB-CO | HF | Solved by resetting the crates. Federico called. He checked the HF trigger cables that were OK. He reset the 3 VME crates. Them the test was repeated successfully. Federico advised to : Quit DAQ in TEMA Rest DAQ crates. Launch HF DAQ in local server. |
| Saturday_22_M | TBF | TEMA completely blocked. | AB-CO | TEMA | TEMA closed. Sever closed and re-launched but no success: problem to be solved Monday by specialists. |
| Saturday_22_M | TBD2 | SPA could not be launched: Power application hung. | AB-CO | Power | Power quitted and re-started: OK |
| Saturday_22_A | TBC1 | No valid data for diode test in HF file. | MTM |  | Federico came and discovered that 2 cables between the front outpout of the potaim cards and the 1st DAQ crate was missing. |
| Saturday_22_A | TBA2 | Special LHC cycle: Communication errors coming from MMP when doing auto test | AB-CO | MMP |  |