Vc/9-Aug-04

### Week32 starting 2Aug  04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 02 aug04  18.15 hrs

| TBA1 | MBAR1139 | ICS 3 Traveller after warm test | 303.6 |
| --- | --- | --- | --- |
| TBA2 | MBBL3191 | Prep 5 Pumping + Cool Down | 90.8 |
| TBB1 | NONE | Bench modif to Spec SSS | 302.5 |
| TBB2 | MBBL2090 | PTE Min. NRJ 3KA Diode Test | 1.9 |
| TBC1 | MBAL3172 | PT 12 Warm Up | 301.9 |
| TBC2 | NONE | Bench modif to Dipole | 302.5 |
| TBD1 | SSS023 | ICS 5 Traveller Last | 302.5 |
| TBD2 | SSS036 | Stretched Wire Warm SW11 | 301.4 |
| TBE1 | MBBL3137 | PT 10 Min. NRJ 11850 A | 11.6 |
| TBE2 | _ | _ | 303.1 |
| TBF1 | MBAR1140 | PT 6.9 Training 9 | 2.35 |
| TBF2 | MBBL1119 | PT 2 HV Cold Test | 1.9 |

8/2/2004 6:11:33 PM    Temperature data are taken from TBSM18 site

Coordinator this week: Stephane ; MTM support on call :  (16-3602).

10 Magnets( 8 dipoles+ 2 SSS)   Disconnected after tests last week    ( Sat Midnight to Sat midnight)  ; + 3 Dipoles & 1 SSS that finished Training & warm up launched  but not met Sat midnight cutoff [ A1: 3177, E2: 2035 , B1:3190 , D1 : SS023 ]

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.(Jerome ?)

    - Modifs to C2 to take on Dipoles

    - Modifs to B1 to take on SSS & later Special SSS’s

  - SSS work :

    - To Do Lists & SMTMS updates & simplifications

    - Diagnostics & Faults chasing

    - Results & “goodness” evaluation

Bench status  and tests for now( Monday 2 Aug/7at 18.20 hrs)  & next few days :  see Video Display

| Magnets under Test on 02-08 6:17:33 PM |
| --- |
| - | Bench | Magnet | Shafts | Start | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | Last Test Desc | Time | I | Temp. | Total |
| 1 | TBA1 | MBAR1139 |  | 02-08 | 14 h. | - | - | - | - | ICS 3 Traveller after warm test | 02-08 (16 : 00) | 0 A. | 300 K | 14 hours |
| 2 | TBA2 | MBBL3191 |  | 31-07 | 34 h. | 21 h. | - | - | - | Prep 5 Pumping + Cool Down | 01-08 (21 : 34) | 0 A. | Cooling | 55 hours |
| 3 | TBB1 | NONE |  | 02-08 | 8 h. | - | - | - | - | Bench modif to Spec SSS | 02-08 (10 : 18) | 0 A. | 300 K | 8 hours |
| 4 | TBB2 | MBBL2090 |  | 29-07 | 63 h. | 19 h. | 26 h. | - | - | PTE Min. NRJ 3KA Diode Test | 02-08 (04 : 37) | 3000 A. | 1.9 K | 108 hours |
| 5 | TBC1 | MBAL3172 |  | 26-07 | 48 h. | 71 h. | 34 h. | 22 h. | - | PT 12 Warm Up | 02-08 (06 : 00) | 0 A. | 4.4 K | 175 hours |
| 6 | TBC2 | NONE |  | 02-08 | 8 h. | - | - | - | - | Bench modif to Dipole | 02-08 (10 : 11) | 0 A. | 300 K | 8 hours |
| 7 | TBD1 | SSS023 |  | 19-07 | 98 h. | -84 h. | 298 h. | 17 h. | 1 h. | ICS 5 Traveller Last | 02-08 (08 : 50) | 0 A. | 300 K | 340 hours |
| 8 | TBD2 | SSS036 |  | 28-07 | 120 h. | - | - | - | - | Stretched Wire Warm SW11 | 02-08 (10 : 03) | 0 A. | 300 K | 120 hours |
| 9 | TBE1 | MBBL3137 |  | 28-07 | 58 h. | 29 h. | 43 h. | - | - | PT 10 Min. NRJ 11850 A | 02-08 (13 : 40) | 11850 A. | 1.9 K | 130 hours |
| 10 | TBF1 | MBAR1140 |  | 27-07 | 69 h. | 2 h. | 74 h. | - | - | PT 9 CDAP HV | 02-08 (17 : 00) | 0 A. | 1.9 K | 145 hours |
| 11 | TBF2 | MBBL1119 |  | 29-07 | 33 h. | 61 h. | - | - | - | PT 2 HV Cold Test | 02-08 (10 : 30) | 0 A. | 4.4 K | 94 hours |
| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| 1 | MBAR1139 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBL3191 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 3 | NONE |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 4 | MBBL2090 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 5 | MBAL3172 | 11656 Amp. (D2_U) | 12851 Amp. () | - | - | - | - | - | - | - | - | - | - | - |
| 6 | NONE |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS023 | 12287 Amp. (QI-34) | 12616 Amp. (QI-12) | 12686 Amp. (QE_12) | 12845 Amp. (QE_12) | 12855 Amp. (QE_34) |  | - | - | - | - | - | - | - |
| 8 | SSS036 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 9 | MBBL3137 | 11089 Amp. (D2-L) | 12296 Amp. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |
| 10 | MBAR1140 | 11202 Amp. (D2-L) | 11645 Amp. (D2-L) | 12142 Amp. (D2-L) | 12383 Amp. (D1-L) | 12513 Amp. (D1-U) | 12597 Amp. (D2 L) | 12662 Amp. (D1-U) | 12700 Amp. (D2-U) | 12332 Amp. (D2_U) |  | - | - | - |
| 11 | MBBL1119 |  | - | - | - | - | - | - | - | - | - | - | - | - |

# Results by Week for PREVIOUS week

| Magnets under Test from 25-07-2004 to 31-07-2004 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBAR1136 |  | 14-07-2004 | 42 h. | 34 h. | 151 h. | 17 h. | 11 h. | 25-07-2004 | 264 hours |  |  |
| 2 | TBA2 | MBBL3189 |  | 24-07-2004 | 46 h. | 43 h. | 65 h. | 16 h. | 9 h. | 31-07-2004 | 168 hours |  |  |
| 3 | TBB1 | MBBL1098 |  | 24-07-2004 | 39 h. | - | - | - | - | 26-07-2004 | 48 hours |  |  |
| 4 | TBB2 | MBAL3180 |  | 22-07-2004 | 28 h. | 55 h. | 32 h. | 20 h. | 3 h. | 29-07-2004 | 168 hours |  |  |
| 5 | TBC1 | MBAL3174 |  | 19-07-2004 | 29 h. | 60 h. | 35 h. | 12 h. | 23 h. | 26-07-2004 | 168 hours |  |  |
| 6 | TBC2 | SSS031 |  | 22-07-2004 | 69 h. | 24 h. | 107 h. | 10 h. | 5 h. | 31-07-2004 | 216 hours |  |  |
| 7 | TBD2 | SSS017 |  | 02-07-2004 | 178 h. | 147 h. | 270 h. | 15 h. | 14 h. | 28-07-2004 | 624 hours |  |  |
| 8 | TBE1 | MBBL2091 |  | 21-07-2004 | 16 h. | 88 h. | 21 h. | 19 h. | 25 h. | 28-07-2004 | 168 hours |  |  |
| 9 | TBF1 | MBBL1129 |  | 19-07-2004 | 35 h. | 66 h. | 64 h. | 23 h. | 1 h. | 27-07-2004 | 192 hours |  |  |
| 10 | TBF2 | MBBL2085 |  | 21-07-2004 | 8 h. | 8 h. | 142 h. | 18 h. | 1 h. | 29-07-2004 | 192 hours |  |  |

+ 3 Dipoles & 1 SSS that finished Training & warm up launched  but not met Sat midnight cutoff [ A1: 3177, E2: 2035 , B1:3190 , D1 : SS023 ]

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR1136 | 10853 A. (D2-L-S13) | 11213 A. (D2-L-S12) | 11451 A. (D2-U S13) | 11722 A. (D2-L-S13) | 11892 A. (D2_U_S13) | 12037 A. (D1_U_S01) | 12193 A. (D1_L_S13) | 12244 A. (D2_L_S01) | 12371 A. (D2-L-S12) | 12036 A. (D2-L-S13) | 12363 A. (D1-LS 13) |  | - |
| 2 | MBBL3189 | 11139 A. (D2-L) | 12270 A. (D2 U) |  | - | - | - | - | - | - | - | - | - | - |
| 3 | MBBL1098 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 4 | MBAL3180 | 8753 A. (D1-L) | 12006 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 5 | MBAL3174 | 10739 A. (D1_U) | 11885 A. (D1_U) | 12300 A. (d2_L) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 6 | SSS031 | 12679 A. (D1_L) | 12823 A. (Q-E-12) | 12851 A. | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS017 | 11325 A. (Q_E 12) | 11607 A. (QI_12) | 12095 A. (QI-12) | 12541 A. (QI-34) | 11431 A. (No Data) | 12850 A. | - | - | - | - | - | - | - |
| 8 | MBBL2091 | 11983 A. (D2_U) | 12859 A. | - | - | - | - | - | - | - | - | - | - | - |
| 9 | MBBL1129 | 12217 A. (D2_L-S01) | 12589 A. (D2_U_s13) |  | - | - | - | - | - | - | - | - | - | - |
| 10 | MBBL2085 | 11527 A. (D1-U-S01) | 12533 A. (D1-U-S01) |  | - | - | - | - | - | - | - | - | - | - |

General Problems last week :

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _27_A Wednesday_28_A | TBC2 | Power supply on Auxiliary circuit cannot be started. | MTM |  | Power supply tested by AB PO in short circuit. OK L measured on circuit CD and EF, found inductances of 5 and 10 mH instead of 18 mH and 62 mH. |
| Tuesday _27_A | TBD1 | HV Test : Report incomplete and test sequence on SS23 did not match with that used for previous SSS (SSS17). | MTM/AB/CO | HV | Federico and TC informed. Richard: Sequence of the dipole was launched by mistake. Then aborted manually and switched to SSS sequence. But generation of the report perturbed. It takes all the files to generate the report. |
| Wednesday _28_A Thurday_29_N | TBA2 | 3189 dipole HV test failed “dipole-QH- vs ground” at 3050 V Happened again at the end of the test program | MTM | HV | Repeated by OP team and by Richard: passed at the end. A.Siemko and Richard decided to give the OK for the magnet. |
| Wednesday_28_N | TBD1 | Vtap VT4 connected to the lead F is opened so the potaim-card Vclamp F is triggering. Vtaps not properly soldered? | ICS? |  | Solved by Federico:bridge in the output of the card done by Federico between VT4 and VT3. |
| Wednesday_28_N Thursday_29_A | TBD1 | Cryo OK cannot be obtained. | ACR |  | Temp problem on auxiliary leads. Potentiometer installed to keep the valve opened. But still not working. Lionel Herblin called by Federico. Manual mode needs more parameter than the automatic mode. Problem solved. |
| Thursday _29_M | TBA2 | SPA happened at 6500 A during the training | ACR-AB/CO |  | Origin not known: Communication problem between PLC and CTH? |
| Thursday _29_A | TBC | Water problem. |  |  | Other valves closed and then interlock disappeared. Origin not found. NB: Bench consumption: 80 l/min. Total available: 440 l/min.: So 4 benches could work in parallel in theory. |
| Thursday _29_A | TBE | IAP : Offset on Potaim card D2_U-L at 5 V: | MTM |  | Federico came and changed the card. |
| Friday_30_A | TBD1 | SSW measurements at 5000 A : During ramping down from 5000 A to 150 A, a water interlock appeared. SPA abort at 5000 A during SSW. | ACR-AB/CO |  | Water closed for other benches. Origin not found. Test repeated and then OK. |
| Friday_30_N | TBF1 | Water fault on PC during quench at 1.5 kA. |  |  | Water flow kept only on two benches. |
| Saturday_31_M Satrurday_31_A Saturday _31_N | TBD1 | Gdl measurements with SSW on SSS23 (on total 11 times): -wire broke 4 times. -SPA abort appeared on PC (2 times) -After the measurements the wire was found at 9mm below the centre in stage B : signal not good (4 times) Wire centred but TF value two low (signal too low). | MTM ACR-AB/CO |  | Anticryostat problem? Wire touching the anticryostat. Return wire not connected to stage B Diagnosis: R between wire and anticryostat : ~0.5 MOhm. |
| Saturday_31_M | TBC | IAP at cold : Problem with the intercon box. Switch not possible. | MTM |  | Federico investigated the problem. Signal coming to the PLCSE not correct because the positioning of the auxiliary switch was not correct. |
| Saturday_31_N | TBE2 | Magnet 2035 : Min NRJ for warm-up: AQA error D1_U peak voltage at 150 V. | MTM | AQA | Inform TC. |
| Saturday_31_N Sunday_01_A | TBD2 | IAP at warm on SSS36: No data for Vtaps. | MTM | IAP | Not the same channel name as compared to SSS23. Hf and lf files did not exist for this magnet. Federico created the config files. |
| Sunday_01_A | TBC1 | Diode test : AQA alarm : Miits above the upper limits (23 at 3 kA) | MTM | AQA |  |
| Sunday_01_A | TBD1 | Slow PA during ramp for fast de-excitation test. | ACR-AB/CO |  | Origin unknown |