Vc/29-Nov-04

### Week49 starting 29 Nov 04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 29 Nov 04  11/29/2004 3:39:23 PM

LHC SM18 Magnet Test Facility

| 4 | A1 | MBBR3017 | PT 12 Warm Up | 212.4 |
| --- | --- | --- | --- | --- |
| 11 | A2 | MBBL3238 | Prep 0 Scan Bench | 294.6 |
| 12 | B1 | _ | _ | 294.7 |
| 9 | B2 | MBAL2120 | Prep 3 CDW1 HV | 296.3 |
| 3 | C1 | MBBR3239 | Prep 5 Pumping + Cool Down | 1.91 |
| 10 | C2 | MBAL3231 | Finished Ready to go Final scan | 296.9 |
| 1 | D1 | SSS030 | Stretched Wire Cold SC25 | 1.9 |
| 8 | D2 | SSS040 | ICS 4 Final connection | 295.5 |
| 2 | E1 | MBAL2122 | PT 6.1 Training 1 | 1.89 |
| 5 | E2 | MBBL1170 | Prep 5 Pumping + Cool Down | 103.6 |
| 6 | F1 | MBBL2118 | PT 12 Warm Up | 62.4 |
| 7 | F2 | MBBR3240 | ICS 4 Final connection | 294.5 |

Temperature data are taken from TBSM18 site

Coordinator this week: Pierre; MTM support on call :   16-3602 .

12 named  Magnets Tested or 13 Magnet  tests  cycles ( leak on ss614  counted as useful test of cryo integrity):

10 dipoles( 2 with MM)+ 2 SSS( 0 with MM ) +  1  SSSS     Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- SSS Problems ( related to SUN ?? )

1) Sometimes the test files are not generated.

2) Sometimes the file is lost while copying.

3) Sometimes during X-Y measurement, X-results are OK But Y-results are exactly "ZERO" OR vise-versa.

4) Some timeswhile copying the result file, PC gives message as "not sufficient space"

- SHUTDOWN for Water : \

- Cluster A and Current leads /water leak repair

- BENCH B1 ( 1 week for Q9 mods gone);  Q9  SSSS  tests start?

- AOB

| Magnets under Test from 21-11-2004 to 27-11-2004 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBBL3219 |  | 20-11-2004 | 33 h. | 26 h. | 43 h. | 16 h. | 9 h. | 25-11-2004 | 120 hours |  |  |
| 2 | TBB1 | SSS614 |  | 18-11-2004 | 34 h. | - | - | - | - | 22-11-2004 | 96 hours |  |  |
| 3 | TBB2 | MBBL2043 |  | 16-11-2004 | 37 h. | 38 h. | 96 h. | 19 h. | 5 h. | 24-11-2004 | 192 hours |  |  |
| 4 | TBC1 | MBAR3247 |  | 23-11-2004 | 16 h. | 21 h. | 17 h. | 22 h. | 10 h. | 27-11-2004 | 96 hours |  |  |
| 5 | TBC1 | MBAR3257 |  | 18-11-2004 | 24 h. | 48 h. | 18 h. | 17 h. | 7 h. | 23-11-2004 | 120 hours |  |  |
| 6 | TBC2 | MBBL1168 |  | 19-11-2004 | 24 h. | 46 h. | 25 h. | 15 h. | 9 h. | 24-11-2004 | 120 hours |  |  |
| 7 | TBD1 | SSS039 |  | 12-11-2004 | 73 h. | 25 h. | 101 h. | 11 h. | 26 h. | 22-11-2004 | 240 hours |  |  |
| 8 | TBD2 | SSS101 |  | 17-11-2004 | 55 h. | 47 h. | 51 h. | 9 h. | 38 h. | 26-11-2004 | 216 hours |  |  |
| 9 | TBE1 | MBBR3259 |  | 20-11-2004 | 34 h. | 27 h. | 47 h. | 30 h. | 7 h. | 26-11-2004 | 144 hours |  |  |
| 10 | TBE2 | MBBL1167 |  | 22-11-2004 | 26 h. | 47 h. | 13 h. | 29 h. | 7 h. | 27-11-2004 | 120 hours |  |  |
| 11 | TBE2 | MBBL3237 |  | 16-11-2004 | 22 h. | 50 h. | 14 h. | 15 h. | 29 h. | 22-11-2004 | 144 hours |  |  |
| 12 | TBF1 | MBBL3216 |  | 20-11-2004 | 33 h. | 44 h. | 28 h. | 21 h. | 7 h. | 26-11-2004 | 144 hours |  |  |
| 13 | TBF2 | MBAL3225 |  | 19-11-2004 | 21 h. | 22 h. | 36 h. | 17 h. | 5 h. | 23-11-2004 | 96 hours |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL3219 | 12661 A. (D1-L-CFB) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 2 | SSS614 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 3 | MBBL2043 | 12049.7 A. (D2_L_S01) | 12529 A. (D1-U-S13) | 12787 A. (D1-U) |  | - | - | - | - | - | - | - | - | - |
| 4 | MBAR3247 | 11625 A. (D2_U) | 12144 A. (D1_U) |  | - | - | - | - | - | - | - | - | - | - |
| 5 | MBAR3257 | 12419 A. (D2 - L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 6 | MBBL1168 | 10815 A. (D2-L) | 12154 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS039 | 11768 A. (QE-34) | 11750 A. (QE-34) | 12755 A. (QE-34) | 12629 A. (QI_12) | 12848 A. (QI-34-S3) | 12851 A. (No Data) | 12851 A. | - | - | - | - | - | - |
| 8 | SSS101 | 11416 A. (QE_34) | 12228 A. (QI-12) |  | - | - | - | - | - | - | - | - | - | - |
| 9 | MBBR3259 | 11710 A. (D2_U) | 11769 A. (D1_U) | 12159 A. (D1-L) | 12298 A. (D1-L) | 12403 A. (D1-L) | 12446 A. (D1_L) | 12593 A. (D1_L) | 12729 A. (D1-L) | 12810 A. (D2_U) |  | - | - | - |
| 10 | MBBL1167 | 12504.7 A. (D1-L) | 12771.6 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 11 | MBBL3237 | 12225 A. (D2-U) | 12659 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 12 | MBBL3216 | 8807 A. (D1-U) | 10518 A. (D2_L) | 12189 A. (D2_L) | 12279 A. (D2_U) | 12484 A. (D2_U) | 12783 A. (D2_U) | 12848 A. | - | - | - | - | - | - |
| 13 | MBAL3225 | 9111 A. (D2-L) | 11183 A. (D1-U) | 12239 A. (D2_U) | 12839 A. (D2 - U) | 12778 A. (D2-L) | 12796 A. (D2-L) | 12850 A. | - | - | - | - | - | - |

| From 21-11-2004 to 27-11-2004 |
| --- |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBF2 | MBAL3225 | False | 19-11-2004 | 21 h. | 22 h. | 36 h. | 17 h. | 5 h. | 23-11-2004 | 4 days | 6 |
| 2 | TBF1 | MBBL3216 | False | 20-11-2004 | 33 h. | 44 h. | 28 h. | 21 h. | 7 h. | 26-11-2004 | 6 days | 6 |
| 3 | TBE2 | MBBL1167 | False | 22-11-2004 | 26 h. | 47 h. | 13 h. | 29 h. | 7 h. | 27-11-2004 | 5 days | 2 |
| 4 | TBE2 | MBBL3237 | False | 16-11-2004 | 22 h. | 50 h. | 14 h. | 15 h. | 29 h. | 22-11-2004 | 6 days | 2 |
| 5 | TBE1 | MBBR3259 | False | 20-11-2004 | 34 h. | 27 h. | 47 h. | 30 h. | 7 h. | 26-11-2004 | 6 days | 9 |
| 6 | TBD2 | SSS101 | False | 17-11-2004 | 55 h. | 47 h. | 51 h. | 9 h. | 38 h. | 26-11-2004 | 9 days | 2 |
| 7 | TBD1 | SSS039 | False | 12-11-2004 | 73 h. | 25 h. | 101 h. | 11 h. | 26 h. | 22-11-2004 | 10 days | 6 |
| 8 | TBC2 | MBBL1168 | False | 19-11-2004 | 24 h. | 46 h. | 25 h. | 15 h. | 9 h. | 24-11-2004 | 5 days | 2 |
| 9 | TBC1 | MBAR3247 | False | 23-11-2004 | 16 h. | 21 h. | 17 h. | 22 h. | 10 h. | 27-11-2004 | 4 days | 2 |
| 10 | TBC1 | MBAR3257 | False | 18-11-2004 | 24 h. | 48 h. | 18 h. | 17 h. | 7 h. | 23-11-2004 | 5 days | 1 |
| 11 | TBB2 | MBBL2043 | True | 16-11-2004 | 37 h. | 38 h. | 96 h. | 19 h. | 5 h. | 24-11-2004 | 8 days | 3 |
| 12 | TBB1 | SSS614 | False | 18-11-2004 | 34 h. | 230 h. | -170 h. | 0 h. | 1 h. | 22-11-2004 | 4 days | 0 |
| 13 | TBA1 | MBBL3219 | True | 20-11-2004 | 33 h. | 26 h. | 43 h. | 16 h. | 9 h. | 25-11-2004 | 5 days | 1 |

from 21-11-2004 to 27-11-2004

2043 had many SPECIAL tests; SS039 on D1 with Shafts; 33219 was 3rd time , now with QAntenna

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _23_A | TBD2 | Problem with cryo OK: PLC communication problem | ACR/MTM |  |  |
| Wednesday _24_A | TBD1 | Launch of HV test at warm failed (Before CT) | AB/CO | HV | HV application re-started then OK |
| Wednesday _24_N Friday _26_A | TBA2 | Water leak coming from power cable. | MTM |  | One leak repaired. Another one appeared. Test finished like this: intervention planned next Tuesday. |
| Wednesday _24_N | TBA2 | Polarity switch not coming in right position | MTM |  | Federico called. He adjusted the position of the switch. |
| Friday_26_N | TBC1 | Too low resistance value of the MCO corrector in the 3247 magnet.: 2.94 Ohms instead of 4.8 Ohms. | MTM |  | NC opened. Magnet put in stand by |
| Friday_26_N | TBC2 | Incorrect name of the magnet in the e-traveller: 3132 instead of 3231. | ICS |  | Name corrected. |
| Friday_26_N | All | Stop of pumping: problem with the BP. | ACR |  | Stop during 6 h30 |
| Saturday 27_M | TBC2 | HV test stopped during execution to permit the removal of the 3247 magnet |  |  | Test re-started after. |
| Saturday 27_A | TBB2 | After final HV magnet warmed up above 10 K. But no quench because no powering. | ACR |  | Because pumping was stopped and the magnet was warming up. |
| Saturday 27_A | TBD2 | Wrong value of temperature on SSS40. | MTM |  | Federico came and changed the setting (calibration). |
| Saturday 27_N | TBD1 | De-excitation test on SSS : cannot impose the value of 350 A for the ramp rate. | MTM | Power |  |
| Sunday 28_M | TBF2 | Leica used for a special test impeded to remove the 1169 magnet | MTM |  | TC informed. He asked to wait for the TS/SU team to see if a measurements is running. |
| Sunday 28_M | TBD1 | MQ strength with SSW repeated fourth times. wire broken, SPA File not generated Results of files lost while copying . Roll angle measurements not OK. Incoherent result from X-Y measurements. | MTM |  | The fourth times OK. Measurement repeated. |
| Sunday 28_A | TBD1 TBF1 | SPA at 5 kA. SPA at 3 kA (2 times) |  |  |  |
| Sunday 28_A | TBA1 | Magnet 3017 measured with Long shaft. Cable connection has to be changed from quench antenna to long shaft. | MTM |  | Federico called. He changed the cable connection. |