### Week 28

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Bench status  and tests for this week :

From Saturday 26/06-3/07 : 9 dipoles+ 1SSS has left the test bench.

- Magnets with Anticryostat: 2065 (special MM test), 2048 (repeated )

Test Coordinator this week:  G.De Rijk;  MTM  Tech. on call Richard:  (16-3602) .

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- This week : more than 5 magnets in liquid phase- priority problems.

- Lost of cryo OK during MM (Saturday and Sunday)

- Magnets non conformties/ and or problems

  - 3143 (E2): YT122 circuit shorted to ground and dipole : detected at warm

  - 3132 (F1) : High peak voltage detected on the first quench : D2, D2_L, >250 V.HF protection circuits more efficient in AP2.

  - 2077 (B1): Step HV  Magnet/vs ground failed at 3.1 kV. IFS box problem.

  - 1069 (F1) : At the first training MIITS too high (39). High RRR?

  - SSS: Leads of EF circuit not isolated (SSS16),  wrong connection of octupole (SSS017) axis meas. for the sextupole is not working with SSSW2…

- AOB

  - New PTOFF version (AQA, hf off, lf off application) on sun MTA 9, 24,25,27,30,31,33,35,37,39

  - Old remains active on sunias08, sun MTA 22,23,42,44.

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_30_M | TBC2 | A trigger was always present on lf acquisition. | MTM | lf | Investigation performed by Maryline She found mistakes in the new config file of the recording of the QH discharge concerning the trigger definition. |
| Tuesday _30_N | TBF1 | 3132 magnet : Peak voltage exceeds the limit for D1, D2, D1_L, D_U, D2_U, D2_L (250 V) | MTM | AQA | Training stopped. Investigation done by P.Pugnat. Origin was a bad coupling for one pair of the heaters. OK to continue given by Pierre, |
| Wednesday _30_M | TBB1 | HV test at cold on 2077: failed first step at 3.1 kV | ICS |  | Special investigation launched. Dipole vs ground failed. IFS box dismounted and replaced. Then OK. |
| Wednesday_31_N | TBE2 | 3143 magnet : IAP at warm shows error for LF discharge. HV test failed. YT122 shorted to the ground. |  |  | Magnet non conform :disconnected. |
| Wednesday_31_N | TBA1 | 2042 magnet : Alarm on AQA; Tau jump too high in ½ HF quench heater. | MTM | AQA | Investigation by Pierre. Doubling of points and OK given by M.Buzio |
| Thursday _01_M | TBA1 | Diode test (PtE) repeated because connector “left” was put on “right” magnet. | MTM, |  | Test repeated without connector. |
| Thursday _01_M | TBC2 | SSS16: Problem with SSW measurement on the sextupole. The option “rotating wire “is not working with SSW Nº 2. Not possible to power the octupole (circuit EF). | MTM |  | Axis from sextupole not measured. Waiting for an answer of J.Di Marco. Investigation done by Salem at warm :inversion of the Vtaps cable+one lead of the EF circuit poorly isolated. |
| Thursday _01_N Sunday_04_N | TBA2 | 1023 magnet : Signal missing for Quench Antenna. 1023: Quench number 9: No quench location data | MTM |  | Guy informed. LQA setting has to be selected. ND: two settings are present in TEMA: QLA right and left depending on the polarity of the magnet. |
| Saturday_03_A Sunday_04_M | TBB1 | LHC cycles on 2077: cryo OK lost 2 times. Too many magnets between1.9 K and 4.2 K? Priority problem? | ACR/MTM |  | NB: LHC cycle have to be repeated after a quench. |
| Saturday_03_N | TBD2 | IAP at warm on SSS017: file stored in the folder of SSS037 because magnet name in TEMA was not changed. | MTM |  |  |
| Sunday_04_M | TBA1 | Problem in signing the e-traveller for the 2048 magnet (repeat magnet) because the steps were signed for the first test. | MTM | TTS | Traveller signed manually. |
| Monday_05_N | TBD1 | SSS20. Not possible to power the EF corrector circuit. | MTM |  | Inductance loop have to be changed from MQT to MO corrector. |
| Monday_05_N | TBD2 | SSS17: Rvalues at warm not correct. | MTM |  | Wrong connection: one sextupole Vtaps not soldered., problem with one lead of EF circuit. |
| Monday_05_N | TBF1 | 1069 magnet : MIITS too high (39). | MTM |  | Training stopped. P.Pugnat informed. RRR measurements decided? |