Vc/12-Apr-05

### Week 15  starting 11 April 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  11 April 05 @ 4/11/2005 3:40:20 PM

| 5 | A1 | MBBL1213 | PTE Min. NRJ 3KA Diode Test | 4.22 HV |
| --- | --- | --- | --- | --- |
| 8 | A2 | MBAL1205 | ICS 4 Final connection | 292.94 |
| 11 | B1 |  |  | 294.55 |
| 10 | B2 | SSS010 | Prep 3 CDW1 HV | 293.83 |
| 1 | C1 | MBBL3636 | PT 6.7 Training 7 | 3.80 HV |
| 7 | C2 | MBAL2168 | ICS 4 Final connection | 293.57 |
| 9 | D1 | SSS151 | ICS 2 Connect Magnet (ICS) | 292.85 |
| 4 | D2 | SSS095 | PT 6.1 Training 1 | 4.06 HV |
| 3 | E1 | MBBR3318 | PT 9 CDAP HV | 1.92 |
| 12 | E2 | MBAL1208 | ICS 4 Final connection | 293.3 |
| 2 | F1 | MBBR3246 | PT 12 Warm Up | 289.64 |
| 6 | F2 | MBAL2153 | Prep 5 Cool Down | 91.64 |

Coordinator this week: John; MTM support on call :   16-3602 .

16  complete Magnet Tests cycles   ( 1193 counted twice)

12  dipoles( 3 with MM)+ 4  SSS(1  with MM) +  0  SSSS (with MM)

Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :

General Points

        - Repeat Magnets Indication in Weekly Reports

        - Strip/stndby indication in Weekly Reports

        - New e-Traveller upgrades for Shafts

The changes only apply to MTM and ICS teams.

For MTM you have 2 new groups of "dual" tasks (you have to sign the first one or the second one) : 
- BEFORE the COLD TESTS: 
716     MTM     SHAFTS INSTALLATION AUTHORIZATION 
717     MTM     NO SHAFTS NEEDED

- AFTER the COLD TESTS: 
721     MTM     SHAFTS REMOVAL AUTHORIZATION 
722     MTM     NO SHAFTS TO REMOVE

For any questions, GHH will be very pleased to help/explain you !

## Tests

| Magnets under Test from 03-04-2005 to 09-04-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | # | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |
| 1 | TBA1 | MBBL1193 |  |  | 29-03-2005 | 3 | 46 h. | 49 h. | 59 h. | 15 h. | 0 h. | 05-04-2005 | 169 h. | Not Found |  |
| 2 | TBA1 | MBBL1193 |  |  | 05-04-2005 | 3 | 2 h. | 21 h. | 26 h. | 16 h. | 11 h. | 08-04-2005 | 76 h. | Stripping |  |
| 3 | TBA2 | MBBL3130 |  |  | 25-03-2005 | 1 | 27 h. | 24 h. | 171 h. | 18 h. | 9 h. | 04-04-2005 | 249 h. | Stripping |  |
| 4 | TBB1 | SSS026 |  |  | 30-03-2005 | 1 | 29 h. | 17 h. | 21 h. | 8 h. | 16 h. | 03-04-2005 | 91 h. | Stripping |  |
| 5 | TBB1 | SSS104 |  |  | 03-04-2005 | 1 | 46 h. | 16 h. | 31 h. | 8 h. | 20 h. | 08-04-2005 | 121 h. | Stripping |  |
| 6 | TBB2 | MBBL1211 |  |  | 31-03-2005 | 1 | 31 h. | 41 h. | 28 h. | 17 h. | 9 h. | 05-04-2005 | 126 h. | Stripping |  |
| 7 | TBC1 | MBBL1212 |  |  | 02-04-2005 | 1 | 33 h. | 31 h. | 14 h. | 16 h. | 13 h. | 07-04-2005 | 107 h. | Stripping |  |
| 8 | TBC2 | MBAR3254 |  |  | 30-03-2005 | 1 | 34 h. | 61 h. | 31 h. | 13 h. | 19 h. | 05-04-2005 | 158 h. | Stripping |  |
| 9 | TBD1 | SSS117 |  |  | 29-03-2005 | 1 | 40 h. | 17 h. | 86 h. | 10 h. | 42 h. | 06-04-2005 | 195 h. | Stripping |  |
| 10 | TBD2 | SSS022 |  |  | 31-03-2005 | 1 | 20 h. | 28 h. | 94 h. | 11 h. | 13 h. | 07-04-2005 | 166 h. | Stripping |  |
| 11 | TBE1 | MBBR3309 |  |  | 30-03-2005 | 1 | 35 h. | 29 h. | 33 h. | 14 h. | 14 h. | 04-04-2005 | 125 h. | Stripping |  |
| 12 | TBE1 | MBBR3315 |  |  | 04-04-2005 | 1 | 22 h. | 24 h. | 29 h. | 12 h. | 14 h. | 08-04-2005 | 101 h. | Stripping |  |
| 13 | TBE2 | MBAL1210 |  |  | 02-04-2005 | 1 | 36 h. | 25 h. | 22 h. | 15 h. | 9 h. | 06-04-2005 | 107 h. | Stripping |  |
| 14 | TBF1 | MBAL2136 |  |  | 02-04-2005 | 1 | 29 h. | 39 h. | 35 h. | 13 h. | 7 h. | 08-04-2005 | 123 h. | Stripping |  |
| 15 | TBF2 | MBAL1206 |  |  | 05-04-2005 | 1 | 38 h. | 23 h. | 19 h. | 18 h. | 10 h. | 09-04-2005 | 108 h. | Stripping |  |
| 16 | TBF2 | MBBR3267 |  |  | 31-03-2005 | 1 | 31 h. | 28 h. | 36 h. | 14 h. | 8 h. | 04-04-2005 | 117 h. | Stripping |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL1193 | 11332.9 A. (D2_U) | 12733.4 A. (D1_U) | 12850 A. | 12900 A. | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBL1193 | 12251.3 A. (D2-L) | 12500 A. (D2-U) | 12850 A. (D1-U) |  | - | - | - | - | - | - | - | - | - |  |
| 3 | MBBL3130 | 8757 A. (D2-U) | 10734 A. (D1_U_S01) | 12083 A. (D2-U-S01) | 12271 A. (D2-U-S01) | 12470.9 A. (D1-L-S01) | 12665 A. (D1_U_S01) | 12790 A. (D2_L_S02) | 12850 A. | - | - | - | - | - |  |
| 4 | SSS026 | 11512 A. (QE-12) | 12177 A. (QE-12) |  | - | - | - | - | - | - | - | - | - | - |  |
| 5 | SSS104 | 11879 A. (QE_34) | 11856 A. (QE-34) | 12670.5 A. (QE-34) | 12850 A. | - | - | - | - | - | - | - | - | - |  |
| 6 | MBBL1211 | 11539.8 A. (D1_L) | 12829 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |  |
| 7 | MBBL1212 | 12402 A. (D2-L) | 12725 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |  |
| 8 | MBAR3254 | 12518 A. (D1_L) | 12715 A. (D1_U) |  | - | - | - | - | - | - | - | - | - | - |  |
| 9 | SSS117 | 11265 A. (QE_12) | 12246 A. (QE-12) |  | - | - | - | - | - | - | - | - | - | - |  |
| 10 | SSS022 | 12192 A. (QE-34) | 11714 A. (QE-12) | 11918 A. (QE-12) | 12036 A. (QE-34) | 12048 A. (QE-12) | 12499 A. (QE_12) | 12245 A. (QE_12) | 12559 A. (QE_12) | 12850 A. | - | - | - | - |  |
| 11 | MBBR3309 | 10188 A. (D2 - U) | 11008 A. (D2 - L) | 12560 A. (D2-L) | 12805.7 A. (D1-L) | 12900 A. | - | - | - | - | - | - | - | - |  |
| 12 | MBBR3315 | 11354 A. (D2 - U) | 12183 A. (D2 - U) | 12900 A. | - | - | - | - | - | - | - | - | - | - |  |
| 13 | MBAL1210 | 11729 A. (D2_L) | 12663 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |
| 14 | MBAL2136 | 12568 A. (D2_U_S01) | 12719 A. (D1-U-S13) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |
| 15 | MBAL1206 | 12096 A. (D1-U) | 12241 A. (D2 - U) | 12534 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |  |
| 16 | MBBR3267 | 12620 A. (D1_L_S12) | 12382 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |  |
| From 03-04-2005 to 09-04-2005 |
| - | Bench | Magnet Name | shafts | Start Date . | # | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q | Signed for |
| 1 | TBF2 | MBAL1206 | False | 05-04-2005 | 1 | 38 h. | 23 h. | 19 h. | 18 h. | 10 h. | 09-04-2005 | 108 h. | 3 | Stripping |
| 2 | TBF2 | MBBR3267 | True | 31-03-2005 | 1 | 31 h. | 28 h. | 36 h. | 14 h. | 8 h. | 04-04-2005 | 117 h. | 2 | Stripping |
| 3 | TBF1 | MBAL2136 | True | 02-04-2005 | 1 | 29 h. | 39 h. | 35 h. | 13 h. | 7 h. | 08-04-2005 | 123 h. | 2 | Stripping |
| 4 | TBE2 | MBAL1210 | False | 02-04-2005 | 1 | 36 h. | 25 h. | 22 h. | 15 h. | 9 h. | 06-04-2005 | 107 h. | 2 | Stripping |
| 5 | TBE1 | MBBR3309 | False | 30-03-2005 | 1 | 35 h. | 29 h. | 33 h. | 14 h. | 14 h. | 04-04-2005 | 125 h. | 4 | Stripping |
| 6 | TBE1 | MBBR3315 | False | 04-04-2005 | 1 | 22 h. | 24 h. | 29 h. | 12 h. | 14 h. | 08-04-2005 | 101 h. | 2 | Stripping |
| 7 | TBD2 | SSS022 | False | 31-03-2005 | 1 | 20 h. | 28 h. | 94 h. | 11 h. | 13 h. | 07-04-2005 | 166 h. | 8 | Stripping |
| 8 | TBD1 | SSS117 | False | 29-03-2005 | 1 | 40 h. | 17 h. | 86 h. | 10 h. | 42 h. | 06-04-2005 | 195 h. | 2 | Stripping |
| 9 | TBC2 | MBAR3254 | False | 30-03-2005 | 1 | 34 h. | 61 h. | 31 h. | 13 h. | 19 h. | 05-04-2005 | 158 h. | 2 | Stripping |
| 10 | TBC1 | MBBL1212 | False | 02-04-2005 | 1 | 33 h. | 31 h. | 14 h. | 16 h. | 13 h. | 07-04-2005 | 107 h. | 2 | Stripping |
| 11 | TBB2 | MBBL1211 | False | 31-03-2005 | 1 | 31 h. | 41 h. | 28 h. | 17 h. | 9 h. | 05-04-2005 | 126 h. | 2 | Stripping |
| 12 | TBB1 | SSS026 | False | 30-03-2005 | 1 | 29 h. | 17 h. | 21 h. | 8 h. | 16 h. | 03-04-2005 | 91 h. | 2 | Stripping |
| 13 | TBB1 | SSS104 | False | 03-04-2005 | 1 | 46 h. | 16 h. | 31 h. | 8 h. | 20 h. | 08-04-2005 | 121 h. | 3 | Stripping |
| 14 | TBA2 | MBBL3130 | True | 25-03-2005 | 1 | 27 h. | 24 h. | 171 h. | 18 h. | 9 h. | 04-04-2005 | 249 h. | 7 | Stripping |
| 15 | TBA1 | MBBL1193 | False | 29-03-2005 | 3 | 46 h. | 49 h. | 59 h. | 15 h. | 0 h. | 05-04-2005 | 169 h. | 2 | Not Found |
| 16 | TBA1 | MBBL1193 | False | 05-04-2005 | 3 | 2 h. | 21 h. | 26 h. | 16 h. | 11 h. | 08-04-2005 | 76 h. | 3 | Stripping |

### Week 13 march 3  to 9 April  05

### [   1193 On bench rapid Thermal cycle,   SSS117( SSW + MM )    &   3130 had extended MM tests ]

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _05_A | TBD1 | SSW at warm on SSS117: -Problem on roll angle measurements - Current reading is not working | MTM |  | TC advised to proceed. Guy solved the problem (current not connected properly) |
| Tuesday _05_N | TBF1 | Spikes seen on MB2136 during PTE (diode test) | MTM |  | TC checked and advised to go on with training. |
| Tuesday _05_N | TBB2 | SSS092: connection problem Lyre should be pushed inside the magnet. | CRI |  | TC asked to stop the connection until Wednesday morning and a check. |
| Tuesday _05_N | TBB1 | CD corrector circuit of SSS104 cannot be powered due to trigger on potaims cards (sext Ext-Int & V-clamp-D) Lost of cryo OK during the powering of the corrector During PT6.1, Some Noise spikes observed in YT312 QH Discharge Current signal. However, No Alerts in AQA. | MTM ACR MTM |  | Giorgio informed. |
| Tuesday _05_N | TBF1 | Cryo OK not stable on MB2136 magnet | ACR |  | LHC Cycle delayed by 1 hour. |
| Wednesday_06_A | TBB1 | CD powering. Intermittant temperature interlock on the aux. current leads. No real fault, just the interlock present. | ACR |  |  |
| Wednesday_06_N | all | MTF and web site information are not working, e-mail to edms support has been sent | IT |  | MTF has database problem, Only A1-A2 and B1 were working. |
| Wednesday_06_N | TBF1 | No tools available to remove the shafts. | ICS |  | Tools found in the mechanic office |
| Thursday_07_M | TBA1 | During the ramping of the traini9ng quench number 2 SPA due to cyo loss- | ACR |  |  |
| Thursday_07_N | TBE, TBF | No power on socket Cluster E and F | MTM |  | a differential was switched off. |
| Friday_08_M | TBA1 | Lf DAQ did not arm | MTM | lf | Solved by re-setting the crates. |
| Friday_08_M | TBD2 | Interlock in Vclamp | ICS/MTM |  | OK after re-soldering the cable and changing the potaim card. |
| Friday_08_A Friday_08_N Saturday_09_M | TBC2 | MB3320 :Dipole - QH Vs Ground failed CDAP HV test could not be launched. Error messages are: 1. PLC HV - PLC SE does not allow to launch the test 2. No phase exist for the selected sequence in config file.. | MTM AB/CO | HV | Changed IFS box and PT2 was repeated. Failed again. TC said to proceed further and compare CDAP and CDBP results Giorgio came and found that CDAP cannot be launched .He launched CDBP and renamed the file. |
| Friday_08_A | TBA2 | Voltage tap connector on current lead A is very tight and difficult to remove | MTM |  |  |
| Saturday_09_A | TBD1 | Powering Correctors on SSS102 Trigger on V_ClampC. | ICS |  | Giorgio came and found that Vtap of clamp C is floating (SSS connection not correct). He bridged it. To be detected with IAP@warm.befor cool down |
| Saturday_09_N | TBD1 | The EF corrector of SSS102 is trained. But while ramping down from 600 A to 0, fast PA occurred. Test could be repeated because the power supply cannot be switched on-stand-by. | MTM AB/PO? |  | Giorgio informed who advised to carry out the ramp down at 2A/sec. Giorgio advised to continue with PT8 and PT9 as the corrector was trained. |
| Sunday_10_M | TBB2 | CDBP on SSS133: message indicating that there is already a test file. | AB/CO | HV |  |
| Sunday_10_M | TBC1 | Cannot cool down the magnet below 4.8K. 6 hours since the last quench! | ACR |  |  |
| Sunday_10_N | TBA2 | Cryo-Heater connector (C11) pin broken while putting onto the HV panel for Warm Initial HV test. | ICS |  | Giorgio called and advised to continue with the test without this. He will rectified it tomorrow. |
| Sunday_10_N | TBE2 | Spurious File created E0504110215.qb01 by an unknown trigger. the QH discharged at about 500 Volts. | MTM |  |  |
| Sunday_10_N | TBE2 | Automatic Magnet name from bench selection is not working. Select E2 bench and give E1. | MTM |  | Manually entered the magnet name for the bench. |
| Monday_11_M | TBE1 | Training could not be launched because of Interlock problem: no indication that the QH are charged. | MTM |  | The PLCSE was in safe mode. Intervention of Giorgio, then OK. |