Vc/17-May-05

### Week 17  starting 25 April 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  18 April 05 @ 4/25/2005 2:21:38 PM

| 4 | A1 | MBBR3331 | PT 12 Warm Up | 235.47 |
| --- | --- | --- | --- | --- |
| 7 | A2 | MBBR3293 | PT 13.2 Resist. Meas. | 297.59 |
| 8 | B1 |  |  | 294.73 |
| 5 | B2 | SSS079 | Prep 5 Cool Down | 4.47 HV |
| 1 | C1 | MBBL1127 | PT 9 CDAP HV | 1.90 |
| 11 | C2 | MBBL1105 | PT 13.1 IAP @ Warm | 299.65 |
| 9 | D1 | SSS175 | Prep 0 Scan Bench | 294.55 |
| 10 | D2 |  |  | 294.55 |
| 6 | E1 | MBBR3317 | ICS 4 Final connection | 295.18 |
| 12 | E2 |  |  | 296.97 |
| 2 | F1 | MBBR3310 | PT 12 Warm Up | 11.53 |
| 3 | F2 | MBAL2171 | PT 3 QH Measure | 4.66 HV |

Temperature data are taken from TBSM18 site

Coordinator this week: N Smirnov; MTM support on call :   16-3602

16  complete Magnet Tests cycles   ()

11  dipoles( 1 with MM)+ 5  SSS( 1  with MM) +  0  SSSS (with MM)

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :   J Halik

- General Points

AOB

| Magnets under Test from 17-04-2005 to 23-04-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | # | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |
| 1 | TBA1 | MBBR3316 |  |  | 13-04-2005 | 1 | 24 h. | 27 h. | 37 h. | 12 h. | 22 h. | 18-04-2005 | 122 h. | Stripping |  |
| 2 | TBA1 | MBBR3331 |  |  | 18-04-2005 | 2 | 30 h. | 22 h. | 37 h. | 15 h. | 0 h. | 23-04-2005 | 104 h. | Thermal cycle |  |
| 3 | TBA2 | MBAL2169 |  |  | 15-04-2005 | 1 | 50 h. | 23 h. | 23 h. | 16 h. | 10 h. | 20-04-2005 | 122 h. | Stripping |  |
| 4 | TBB1 | SSS170 |  |  | 15-04-2005 | 1 | 27 h. | 15 h. | 45 h. | 8 h. | 21 h. | 20-04-2005 | 116 h. | Stripping |  |
| 5 | TBB2 | SSS057 |  |  | 20-04-2005 | 1 | 29 h. | 11 h. | 20 h. | 9 h. | 14 h. | 23-04-2005 | 83 h. | Stand By |  |
| 6 | TBB2 | SSS130 |  |  | 15-04-2005 | 1 | 22 h. | 18 h. | 38 h. | 15 h. | 18 h. | 20-04-2005 | 111 h. | Stand By |  |
| 7 | TBC1 | MBBL1111 |  |  | 17-04-2005 | 1 | 32 h. | 35 h. | 25 h. | 14 h. | 11 h. | 21-04-2005 | 117 h. | Stripping |  |
| 8 | TBC2 | MBBL1218 |  |  | 16-04-2005 | 1 | 25 h. | 48 h. | 17 h. | 17 h. | 15 h. | 21-04-2005 | 122 h. | Stripping |  |
| 9 | TBD1 | SSS151 |  |  | 11-04-2005 | 1 | 29 h. | 26 h. | 103 h. | 9 h. | 77 h. | 21-04-2005 | 244 h. | Stripping |  |
| 10 | TBD2 | SSS095 |  |  | 14-04-2005 | 2 | 0 h. | 97 h. | 9 h. | 16 h. | 30 h. | 20-04-2005 | 152 h. | Stripping |  |
| 11 | TBE1 | MBAL1177 |  |  | 12-04-2005 | 1 | 25 h. | 33 h. | 32 h. | 13 h. | 10 h. | 17-04-2005 | 113 h. | Stripping |  |
| 12 | TBE1 | MBBL2164 |  |  | 17-04-2005 | 2 | 28 h. | 27 h. | 37 h. | 13 h. | 0 h. | 22-04-2005 | 105 h. | Thermal cycle |  |
| 13 | TBE2 | MBBL1216 |  |  | 15-04-2005 | 1 | 40 h. | 23 h. | 51 h. | 15 h. | 9 h. | 20-04-2005 | 138 h. | Stripping |  |
| 14 | TBF1 | MBAR3329 |  |  | 17-04-2005 | 1 | 32 h. | 28 h. | 33 h. | 20 h. | 20 h. | 22-04-2005 | 133 h. | Stripping |  |
| 15 | TBF2 | MBAL3162 |  |  | 13-04-2005 | 1 | 25 h. | 26 h. | 26 h. | 13 h. | 10 h. | 18-04-2005 | 100 h. | Stripping |  |
| 16 | TBF2 | MBBR3242 |  |  | 18-04-2005 | 1 | 29 h. | 26 h. | 22 h. | 17 h. | 7 h. | 22-04-2005 | 101 h. | Stripping |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR3316 | 11608 A. (D1-U) | 11727 A. (D2-L) | 11889 A. (D2_L) | 12240 A. (D2 - L) | 12463 A. (D2_L) | 12516 A. (D2-L) | 12669 A. (D2-L) | 12850 A. | - | - | - | - | - |
| 2 | MBBR3331 | 9139 A. (D1_U) | 10656 A. (D1-L) | 11252 A. (D1-U) | 11536.9 A. (D2-U) | 11677 A. (D1_L) | 12094 A. (D1-L) | 12352 A. (D1-L) | 12491 A. (D1_L) | 40 A. (No Data) | 12761 A. (D2_U) |  | - | - |
| 3 | MBAL2169 | 10820 A. (D2-U) | 12137 A. (D2 - U) | 12507 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |
| 4 | SSS170 | 12008 A. (QI_34) | 12266 A. (QE-34) | 12165 A. (QE_34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS057 | 12553.7 A. (QI_34) | 12837.6 A. (QE-34) | 12697 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS130 | 10210 A. (QE_12) | 11361 A. (QE_12) | 12223 A. (QE-12) | 12497 A. (QI-12) | 12780 A. (QI_34) | 12845 A. (QI_12) | 12850 A. | 12850 A. | - | - | - | - | - | - |
| 7 | MBBL1111 | 11781 A. (D2 - L) | 12137 A. (D2 - U) | 12261 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |  |
| 8 | MBBL1218 | 12411 A. (D1-U) | 12840 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |
| 9 | SSS151 | 12389 A. (QI - 12) | 12436 A. (QI-12) | 12832.3 A. (QI-34) | 12850 A. | - | - | - | - | - | - | - | - | - |  |
| 10 | SSS095 | 12630.2 A. (QI-34) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |  |
| 11 | MBAL1177 | 11798 A. (D2 - U) | 12313 A. (D1 - L) | 12856 A. | - | - | - | - | - | - | - | - | - | - |  |
| 12 | MBBL2164 | 10899 A. (D1-U) | 11473 A. (D1-L) | 12179 A. (D1 - U) | 12747 A. (D1 - L) | 12848 A. (D1 - L) | 12443 A. (D1-L) | 12263 A. (D1 - L) | 12240 A. (D1_L) | 11947.7 A. (D1-L) | 11712 A. (D1-L) |  | - | - |  |
| 13 | MBBL1216 | 11359 A. (D2-U) | 11838 A. (D2-U) | 12114 A. (D1 - U) | 12319 A. (D2 - L) | 12400 A. (D2_L) | 12546 A. (D1-U) | 12737 A. (D2-U) | 12841 A. (D2 - L) | 12850 A. | - | - | - | - |  |
| 14 | MBAR3329 | 10744 A. (D1-L-S01) | 12313 A. (D1-L-S01) | 12617 A. (D1-L-S09) |  | - | - | - | - | - | - | - | - | - |  |
| 15 | MBAL3162 | 9081 A. (D2_L) | 11344 A. (D1-U) | 12583 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - | - | - |  |
| 16 | MBBR3242 | 11966 A. (D1-U) | 12013.6 A. (D2-U) | 12133.8 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |  |

### Week 17  :  17 April to 23 April 05  (19 Magnet Tests)

### [   3636 On bench rapid Thermal cycle   ]

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _19_M | TBA1 | Magnet MBBR3331, a NC should be open because bar code on the magnet is written MBAR3331. |  |  |  |
| Tuesday _19_M | TBC2 | MB1216: HV failed between Dipole-GND even with eth IFS box changed. | MTM |  | Federico called. The 1.5 kA quenches and the 3 kA quench were done. Test repeated then OK. |
| Tuesday _19_M | TBA1 | Some flat signals during IAP @ warm | MTM |  | Federico changed a faulty potaim-card. |
| Wednesday_20_M | TBF1 | MM 1 on MB3229.Check by Stephane: polarity error message on one head of aperture 2. | MTM |  | Stephane found that the switch in the amplifier of the coil 0001 was in e position and the setting in a position. MM1 repeated. |
| Wednesday_20_N | TBD2 | IAP@warm. Trigger present on VClampE. | ICS? |  | Federico asked ICS to check the SSS connection. OK after the reconnection of the magnet. |
| Wednesday_20_N | TBD2 | Upgrade of the MTF during the night. Trouble in the HV recognition system Problem remains in the Thursday morning. | TS/CE | MTF | HWREC switched off. Magnet mane entered manually. Giorgio switched off HWREC of TBE and TBB, copied the config files in order to allow OP team to go on with the tests. |
| Wednesday_20_N | TBA1 | Ext2 interlock on RMA. It could not be reset. | AB/PO |  | Federico found a fault inside the discharge switch and has been reseted. A reset of the communication card of RMA was also performed. |
| Thursday_21_N Friday_22_M Sunday _25_M | TBA1 | MBB3331:Faulty quench twice at very low current (40 A approx.). Trigger in Vbus-Ext, Vbus-Int. Idem Idem | AB/PO |  | Federico solved the problem by resetting the communication card on RMA |
| Friday _22_A | TBA1 | WP07 on MB not found | ICS |  | ICS team entered the file result wit the wrong magnet name.. |
| Friday _22_A | all | Server problem : MTF is not working | IT | MTF | Tests performed with HWREC off. |
| Friday _22_A | TBB2 | Warning for temperature sensor discrepancy found during the second training quench of SSS057. | MTM |  | Federico will check the coefficients of the temperature sensor. |
| Friday _22_N | TBB1 | Magnet name error in E-traveller done by ICS SSS96 instead SSS096 | ICS |  |  |
| Friday _22_N | TBB1 | In PT5.2 provoked quench at 1.5 kA, symmetrical quench is detected, but voltage tap signals not reaching 200mV. | MTM | hf | DAQ triggered on a spike. |
| Friday _22_N | TBA2 | During IAP @ cold, not possible to arm HF DAQ. | AB/CO | hf |  |
| Saturday_23_M | TBA2 | During IAP @ cold on MB3293, few offset of QLoc cards were not very stable, a warning appeared AC compensation signals not found | MTM | IAP | Federico checked and advised to proceed with the tests. P20 connector not connected to the IFS box |
| Saturday_23_A | TBC2 | Cryo problem on MBBL1105? | ACR |  |  |
| Saturday_23_N | TBA2 | MBB3293: errors on MM1 no temperature value no cmp signals bucking ratio too high After the LHC cycle red alarms on the angle differences in the plateaus. | MTM | MMP | Federico called stephane. MMP settings not done in cmp mode. Temperature problem unsolved. Stephane advised to proceed with the tests Stephane asked to stop the MM. Finish the test and remove the shafts. |
| Saturday_23_N |  | On SSS079 the shock-log box was not removed |  |  | TC asked to remove the box. |
| Sunday_24_M | TBC2 | HV test at cold failed in MBBL1105 QH vs ground, capacitance value out of range for one QH | MTM |  | Federico asked to start the special procedure checking the QH one by one. He phoned to A.Siemko for advice. He advised to go on with the test. |
| Sunday_24_M | TBA2 | Magnet has temperature rise after discharge of QH | MTM |  | Federico contacted Giorgio. OK to continue the test. Check Monday. |
| Sunday_24_N | TBA1 TBE1 | PT8 generated in folder "MBBR_2_2_2_2”. The file for Prep 2.1 (IAP@warm LF) was generated in 3317_2 folder even though it's not repeat magnet. | AB/CO | TEMA | Routine corrected by Marco Monday Morning. |
| Monday_25_M | all | No MTF | TS/CE | MTF | Work with HRWREC off. |