### Week 12 starting 20 march 06

### SM18 Operation Meeting (Tuesdays) 21 March 06

Bench status  on Monday  20-03- 2006 at: 3/20/2006 10:02:31 AM

| 12 | A1 | MBBL1109 | PT 9 CDAP HV | 2.35 |
| --- | --- | --- | --- | --- |
| 5 | A2 | MBBR1284 | ICS 4 Final connection | 299.11 |
| 11 | B1 | SSS504 | PT 7 Magnet Trained | 6.09 |
| 1 | B2 | SSS094 | PT 10.1: Prov. Quench 1/2 HF 11850 | 2.31 HV |
| 10 | C1 | SSS233 | PT 3 QH Measure | 90.19 |
| 9 | C2 | SSS183 | PT 3 QH Measure | 7.10 |
| 6 | D1 | SSS620 | Prep 0 Scan Bench | 297.59 |
| 7 | D2 | SSS611 | PT 14 Magnet MAPS | 299.2 |
| 4 | E1 | MBBR2288 | Prep 5 Cool Down | 96.27 |
| 2 | E2 | MBBR2278 | PT 6.6 Training 6 | 4.58 |
| 3 | F1 | MBBR1279 | PT 8 Fast De-excitation | 1.90 |
| 8 | F2 | MBAL1064 | ICS 1 WP04 HV Test Warm | 299.11 |

Tempr. from TBSM18 site

Tempr. from TBSM18 site

Supervisor: V.Chohan;           Eq support : 16-3602 :

16 (17)  Magnet Tests completed :  8 Dipoles (1214  REPEAT)

: 6 Arc-SSS ( sick 38, orbit dip HV on 236)

: 2 IR-SSS  [637, 621( long due to 4.4 K & anticryo & MM ) ]

General Points :

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- TOO much rush in putting 620 & 504 & total waste of time over w’end

  - 620 pipes too long ..now need insitu cutting ( 72 + hr bench time lost !)

  - ROCLA & Tonio & parking 621

  - IFS box for 620 not in standard list ( ‘not to be tested in SM18” !)???

  - For 504 on B1 :MRB switch for cryo

  - For 504 on B1: jumper in ‘libre’ connector & night shift Fri IAP@warm OP efforts wasted !!

- Situation from (SSS, SSSS)  :  J.Halik

- MANY HV problems currently !!

A.O.B.

## Tests

| Magnets under Test from 12-03-2006 to 18-03-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR2286 |  |  | 11-03-2006 | 26 h. | 29 h. | 20 h. | 16 h. | 10 h. | 15-03-2006 | 101 h. | Stripping |  |  |
| 2 | TBA2 | MBBR1310 |  |  | 11-03-2006 | 39 h. | - | - | - | - | 13-03-2006 | 44 h. | Stand By |  | Leg Leak |
| 3 | TBB1 | SSS038 |  |  | 07-03-2006 | 53 h. | 16 h. | 126 h. | 12 h. | 8 h. | 16-03-2006 | 215 h. | Stand By |  | SICK!! |
| 4 | TBB2 | SSS230 |  |  | 13-03-2006 | 26 h. | 31 h. | 18 h. | 7 h. | 23 h. | 18-03-2006 | 105 h. | Stripping |  |  |
| 5 | TBB2 | SSS248 |  |  | 09-03-2006 | 24 h. | 27 h. | 17 h. | 8 h. | 17 h. | 13-03-2006 | 93 h. | Stripping |  |  |
| 6 | TBC1 | SSS227 |  |  | 09-03-2006 | 23 h. | 21 h. | 18 h. | 6 h. | 16 h. | 13-03-2006 | 84 h. | Stripping |  |  |
| 7 | TBC1 | SSS236 |  |  | 13-03-2006 | 28 h. | 40 h. | 15 h. | 6 h. | 15 h. | 18-03-2006 | 104 h. | Stand By |  | Orbit Dipole HV |
| 8 | TBC2 | SSS325 |  |  | 11-03-2006 | 26 h. | 15 h. | 73 h. | 9 h. | 8 h. | 17-03-2006 | 131 h. | Stripping |  |  |
| 9 | TBD1 | SSS621 |  |  | 08-03-2006 | 60 h. | 13 h. | 144 h. | 8 h. | 20 h. | 18-03-2006 | 245 h. | Stand By |  |  |
| 10 | TBD2 | SSS637 |  |  | 10-03-2006 | 43 h. | -22 h. | 61 h. | 9 h. | 9 h. | 15-03-2006 | 100 h. | Stripping |  |  |
| 11 | TBE1 | MBAL1162 |  |  | 14-03-2006 | 24 h. | 24 h. | 30 h. | 15 h. | 8 h. | 18-03-2006 | 101 h. | Stand By |  | HV |
| 12 | TBE1 | MBAR1325 |  |  | 10-03-2006 | 25 h. | 21 h. | 15 h. | 15 h. | 20 h. | 14-03-2006 | 96 h. | Stripping |  |  |
| 13 | TBE2 | MBAR2245 |  |  | 13-03-2006 | 18 h. | 26 h. | 32 h. | 13 h. | 16 h. | 17-03-2006 | 105 h. | Stand By |  | Detrain |
| 14 | TBE2 | MBBL1214 |  |  | 08-03-2006 | 27 h. | 23 h. | 15 h. | 19 h. | 13 h. | 12-03-2006 | 97 h. | Stripping |  |  |
| 15 | TBF1 | MBAR1324 |  |  | 13-03-2006 | 26 h. | 25 h. | 17 h. | 14 h. | 19 h. | 17-03-2006 | 101 h. | Stripping |  |  |
| 16 | TBF1 | MBBR2256 |  |  | 08-03-2006 | 32 h. | 28 h. | 21 h. | 15 h. | 12 h. | 13-03-2006 | 108 h. | Stand By |  | diode |
| 17 | TBF2 | MBAR1327 |  |  | 11-03-2006 | 17 h. | 27 h. | 20 h. | 13 h. | 8 h. | 15-03-2006 | 85 h. | Stripping |  |  |

### Qu

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR2286 | 10285 A. (D1-U) | 11529 A. (D2-U) | 12228.6 A. (D1-U) | 12614.7 A. (D1-U) | 12852 A. | - | - | - | - | - | - | - | - |
| 2 | MBBR1310 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 3 | SSS038 | 12850 A. | - | - | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS230 | 11908.4 A. (QE_12) | 12416.8 A. (QI-12) | 12840 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS248 | 12343 A. (QE_34) | 12502 A. (QI_12) | 12672.7 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS227 | 11928.6 A. (QI_12) | 12811 A. (QE_12) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS236 | 12476.8 A. (QI-12) | 12676.9 A. (QI-34) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 8 | SSS325 | 10903.5 A. (QI-34) | 11799.8 A. (QI-12) | 11837.8 A. (QI-34) | 12317.1 A. (QI_34) | 12446.1 A. (QI_12) | 12509.1 A. (QI_34) | 12728.2 A. (QI_34) | 12851.5 A. (QE_34) | 12851 A. | - | - | - | - |
| 9 | SSS621 | 3598 A. (MQY-34) | 3681.8 A. (E_MQY_12) | 3717.3 A. (E_MQY_34) | 3761.9 A. (I_MQY_34) | 3787.2 A. (E-MQY-12) | 3776.3 A. (MQY_E_3) | 3811 A. (MQY-32) | 3803.2 A. (MQY-E-32) | 3861.8 A. (E-MQY-32) | 3878.6 A. (I-MQY-12) | 3894.1 A. (E-MQY-12) | 3900 A. | - |
| 10 | SSS637 | 5792 A. (MQM_ext) | 5820 A. | - | - | - | - | - | - | - | - | - | - | - |
| 11 | MBAL1162 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 12 | MBAR1325 | 12332 A. (D2-U) | 12471 A. (D1-L) | 12377 A. (D1-U) |  | - | - | - | - | - | - | - | - | - |
| 13 | MBAR2245 | 10221.3 A. (D1_L) | 11623 A. (d2-L) | 12659.9 A. (D1_U) | 12371 A. (D1-U) | 11713 A. (D1-U) | 11719 A. (D1-U) |  | - | - | - | - | - | - |
| 14 | MBBL1214 | 11963 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 15 | MBAR1324 | 12085 A. (D2_U) | 12557.7 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 16 | MBBR2256 | 10622 A. (D2_L) | 12380 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 17 | MBAR1327 | 11829 A. (D1-L) | 12400 A. (D1-L) | 12419.3 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |

SSS038 sick etc ask stafiniak;    621 had long training due to anti-cryo & MM;      1310 had leak on leg