### Week 17 starting 24 april 06

### SM18 Operation Meeting (Tuesdays) 25 April 06

Bench status  on Monday  24-04-2006 at: 4/24/2006 1:03:44 PM

| 8 | A1 | MBAR2323 | PTE Min. NRJ 3KA Diode Test | 4.11 |
| --- | --- | --- | --- | --- |
| 3 | A2 | MBAR1272 | PT 12 Warm Up | 4.43 |
| 9 | B1 | SSS322 | PT 12 Warm Up | 300.1 |
| 2 | B2 | SSS247 | PT5.6: Powering Corrector IJ | 4.52 HV |
| 11 | C1 |  |  | 2.31 HV |
| 7 | C2 | SSS246 | PT 12 Warm Up | 1.91 |
| 12 | D1 | SSS636 | ICS 4 Final connection | 298.76 |
| 5 | D2 | SSS503 | PT 12 Warm Up | 295.18 |
| 1 | E1 | MBBL2180 | Prep 2.1 IAP @ warm LF | 300.82 |
| 6 | E2 | MBBR1313 | PT 13.2 Resist. Meas. | 300.82 |
| 10 | F1 | MBBR2298 | Prep 5 Cool Down | 95.05 |
| 4 | F2 | MBAR1336 | PT 14 Magnet MAPS | 299.2 |

Tempr. from TBSM18 site

Supervisor: Marco/        Eq support : 16-3602 :

18 Magnet Tests completed :  8 Dipoles (),    7 Arc-SSS  (294  REPEATed)

:		 0 500-series ( )

:							 2 IR-SSS ( 632, 672 (R) )

General Points :

- C Cluster being converted  now for Dipoles Testing (run out of Arc-SSS’s !)

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :OP issues :

- Strike ( of CERN staff) on Friday 28th; public holiday 1st May

- A.O.B.

| Magnets under Test from 16-04-2006 to 22-04-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAR2235 |  |  | 11-04-2006 | 27 h. | 21 h. | 24 h. | 16 h. | 15 h. | 16-04-2006 | 103 h. | Stripping |  |  |
| 2 | TBA1 | MBAR2243 |  |  | 16-04-2006 | 29 h. | 20 h. | 24 h. | 17 h. | 29 h. | 21-04-2006 | 119 h. | Stripping |  |  |
| 3 | TBA2 | MBBR2219 |  |  | 12-04-2006 | 39 h. | 24 h. | 76 h. | 18 h. | 11 h. | 19-04-2006 | 168 h. | Stripping |  |  |
| 4 | TBB1 | SSS231 |  |  | 12-04-2006 | 22 h. | 15 h. | 53 h. | 7 h. | 15 h. | 17-04-2006 | 112 h. | Stripping |  |  |
| 5 | TBB1 | SSS323 |  |  | 17-04-2006 | 29 h. | 13 h. | 17 h. | 9 h. | 26 h. | 21-04-2006 | 94 h. | Stripping |  |  |
| 6 | TBB2 | SSS288 |  |  | 18-04-2006 | 24 h. | -17 h. | 43 h. | 8 h. | 13 h. | 21-04-2006 | 71 h. | Stripping |  |  |
| 7 | TBB2 | SSS332 |  |  | 13-04-2006 | 42 h. | 0 h. | 58 h. | 10 h. | 13 h. | 18-04-2006 | 123 h. | Stripping |  |  |
| 8 | TBC1 | SSS326 |  |  | 14-04-2006 | 37 h. | 15 h. | 52 h. | 10 h. | 15 h. | 20-04-2006 | 129 h. | Stripping |  |  |
| 9 | TBC2 | SSS294 |  |  | 14-04-2006 | 32 h. | 13 h. | 36 h. | 5 h. | 0 h. | 18-04-2006 | 86 h. | Not Found or Thermal cycle |  |  |
| 10 | TBC2 | SSS294 |  |  | 18-04-2006 | 4 h. | 10 h. | 16 h. | 6 h. | 23 h. | 20-04-2006 | 59 h. | Stripping |  |  |
| 11 | TBD1 | SSS672 |  |  | 06-04-2006 | 55 h. | - | - | - | - | 16-04-2006 | 209 h. | Not Found or Thermal cycle |  |  |
| 12 | TBD1 | SSS672 |  |  | 16-04-2006 | 0 h. | 20 h. | 5 h. | 23 h. | 10 h. | 18-04-2006 | 58 h. | Stand By |  |  |
| 13 | TBD2 | SSS632 |  |  | 14-04-2006 | 21 h. | 34 h. | 35 h. | 8 h. | 21 h. | 20-04-2006 | 119 h. | Stand By |  |  |
| 14 | TBE1 | MBAL1091 |  |  | 13-04-2006 | 30 h. | -17 h. | 67 h. | 13 h. | 19 h. | 18-04-2006 | 112 h. | Stripping |  |  |
| 15 | TBE2 | MBAR2228 |  |  | 15-04-2006 | 28 h. | 26 h. | 16 h. | 15 h. | 12 h. | 19-04-2006 | 97 h. | Stripping |  |  |
| 16 | TBF1 | MBAR1267 |  |  | 18-04-2006 | 25 h. | 23 h. | 18 h. | 15 h. | 19 h. | 22-04-2006 | 100 h. | Stripping |  |  |
| 17 | TBF1 | MBAR1304 |  |  | 13-04-2006 | 26 h. | 24 h. | 29 h. | 15 h. | 11 h. | 18-04-2006 | 105 h. | Stripping |  |  |
| 18 | TBF2 | MBBR1330 |  |  | 15-04-2006 | 36 h. | 23 h. | 19 h. | 20 h. | 9 h. | 20-04-2006 | 107 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR2235 | 10557 A. (D2-L) | 10975 A. (D2-L) | 12248 A. (D1-U) | 12773.4 A. (D1_U) | 12853 A. | - | - | - | - | - | - | - | - |
| 2 | MBAR2243 | 11150.2 A. (D1_U) | 12535.5 A. (D1_L) | 12223.5 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 3 | MBBR2219 | 11794.4 A. (D2-L-S13) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS231 | 11048 A. (QI_34) | 11516 A. (QE_12) | 12009.4 A. (QI-12) | 11959 A. (No Data) | 12460.2 A. (QE-34) | 12672.3 A. (QI_34) | 12622.3 A. (QI_34) | 12420.1 A. (No Data) | 12754.8 A. (QE_34) | 12796.8 A. (QE_34) | 12849.4 A. (No Data) | 12850 A. | - |
| 5 | SSS323 | 12286 A. (DI-34) | 12523 A. (QI_12) | 12666.3 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS288 | 12102.4 A. (QI_34) | 12450 A. (QE_12) | 12643.9 A. (QE-12) |  | - | - | - | - | - | - | - | - | - |
| 7 | SSS332 | 11075.9 A. (QE_34) | 12181.1 A. (QI_12) | 12540.4 A. (QE_34) |  | - | - | - | - | - | - | - | - | - |
| 8 | SSS326 | 11961.6 A. (QE-12) | 12366.2 A. (QI-12) | 12444.5 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 9 | SSS294 | 11418 A. (QE_12) | 11734 A. (QE_34) | 11837.9 A. (QE_34) | 11983.3 A. (QE_34) | 12071.7 A. (QE_12) | 12165.2 A. (QE_12) | 12232.8 A. (QE_12) | 12314.7 A. (QE_34) | 12403 A. (QI-34) |  | - | - | - |
| 10 | SSS294 | 12160.2 A. (QE_34) | 12315 A. (QE_12) |  | - | - | - | - | - | - | - | - | - | - |
| 11 | SSS672 | 3290.5 A. (MQY_I_1) | 3554.5 A. (Int_MQY1) | 3733.1 A. (MQY_E_2) | 3737.3 A. (Int_MQY1) | 3688.2 A. (Int_MQY1) | 3733 A. (MQY_I_1) | 3806 A. (No Data) | 3784 A. (No Data) | 3789 A. (E+I_2) | 3798 A. (E+I_2) | 3806 A. (E-MQY-1) | 3825.5 A. (Int_MQY1) | 3827.8 A. (EXt_MQY1) | 3831.5 A. (In_MQY_1) | 3833.5 A. (Ex_MQY_1) | 3838.9 A. (I-MQY-1) | 3861.2 A. (E-MQY-1) | 3901 A. (I-MQY-1) | 3893 A. (Ex_MQY_1) | 3900 A. |
| 12 | SSS672 | 3739.9 A. (Int_MQY2) | 3875 A. (I-MQY-1) | 3879.4 A. (E-MQY-1) | 3900 A. | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |
| 13 | SSS632 | 3780.7 A. (E_MQM_34) | 4255.5 A. (E_MQM_34) | 4345.4 A. (E_MQM_34) | 4644.5 A. (E_MQM_34) | 4650 A. | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |
| 14 | MBAL1091 | 11956.8 A. (D2-U) | 12047 A. (D2-L) | 12668.4 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |
| 15 | MBAR2228 | 11749 A. (D1_L) | 12684.7 A. (D2_L) | 12852 A. | - | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |
| 16 | MBAR1267 | 12000.3 A. (D1_L) | 12197.3 A. (D2_U) | 12744.9 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |
| 17 | MBAR1304 | 12170.8 A. (D1-U) | 12682.8 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |
| 18 | MBBR1330 | 12178.4 A. (D1_L) | 12610.8 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |

S672 over 15 Quenches in 1st run  & had a Th cycle ;           s294 Repeated   so 2 runs   ;     2219 on A cluster  had Magnetic M’ments done