### vc/Week 32 starting 7 aug 06

### SM18 Operation Meeting (Tuesdays) 08 Aug 06

Bench status  on Monday  07-08-2006 at: 8/7/2006 1:22:03 PM

LHC SM18 Magnet Test Facility

| 4 | A1 | MBBR2362 | PT 14 Magnet MAPS | 299.11 |
| --- | --- | --- | --- | --- |
| 12 | A2 | MBAR2333 | ICS 5 Traveller Last | 299.11 |
| 10 | B1 | SSS314 | ICS 4 Final connection | 298.49 |
| 2 | B2 | SSS528 | PT5.3: Powering Corrector CD | 2.35 |
| 11 | C1 | SSS276 | PT 12 Warm Up | 237.58 |
| 5 | C2 | SSS336 | PT 14 Magnet MAPS | 299.02 |
| 6 | D1 | SSS646 | WP17 - HV Test | 297.41 |
| 9 | D2 | SSSS618 | PT 12 Warm Up | 294.64 |
| 1 | E1 | MBAR1303 | ICS 4 Final connection | 299.11 |
| 7 | E2 | MBAL2281 | PT 13.2 Resist. Meas. | 298.04 |
| 3 | F1 | MBAR1301 | PT 3 QH Measure | 2.35 |
| 8 | F2 |  |  | 297.59 |

Major WATER Station cleanout for 30 hrs starting Monday 7thAug 8 am

CONFIG :  D 6 - AS 3 -  IRSSS 3 ;

Supervisor: marco  ;Eq support : 16-3602

14 Magnet Tests completed

6 Dipoles (),

5 Arc-SSS  tests but ONLY 4  magnets( 312 ROBTC)

0  500-series [ ]

2  600-series [ 624 (165 hrs), 650 2nd run – MM done (169 hrs) ]

General Points :

E2 Bench will be  blocked by   2281 for nearly 10-12 days when it goes!!

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :D1 & D2 issues

- A.O.B..

650 ..was 2nd run here  & with MM after problems of anticryo & shafts in 1st run ; 624 had 8 quenches

312 had ROBTC

| 01M (TUE) |
| --- |
| TBA1 | Taking long time to reach 1.9. Checked with CRYO team. They said that they are not sure what is the problem and are looking in to it. | (Software problem - control of valves) |
| TBC | Cluster C : Quench Heater 1 and 2 not charging. Interconn box not switching properly. Frederico informed. He will look in to the problem. | QH charging problem solved by Federico. QH matrix card 1 was not fully inserted. |
| TBE2 | TBE2 : It was observed that a cryo sensor cable was shorting the main magnet current lead on top of CFB. This was sorted out and HV test is re launched. |  |
| “ | TBE2: The HV tests are failing again. waiting again for condensation to clear. |  |
| TBE | Cluster E: Hardware recognition failure. Federico informed. | TT821 coeff data are missing again in MTF database |

| 01A (TUE) |
| --- |
| TBE2 | HV failed again after provoked quench even with a different HV rack. So it is signed for warm up as decided. |  |
| “ | MBAL2281 HV problem persists. Fredrico, in consultation with TC (Arjan) & Pear, decided to perform a provoke quench (1.5kA) & repeat CDBP to check the performance. PT3 was ok. |  |

01 AUG 06 (TUE Morning) to 07 AUG 06 ( MON Night)

| 01N(TUE) |
| --- |
| TBC1 | LF acquisition didn't generate any file during provoked quench (for warming up the magnet for ROBTC), instead the message "Generic error occurred” | ES may please look into this as this is repeatedly coming. |

| 02M(WED) |
| --- |
| TBA1 | The magnet on A1 can not be removed. Rocla access is blocked by 2 vehicles, an anticryostat rack, one SSW stage and black magnet end-caps..... |  |

| 02A(WED) |
| --- |
| TBA1 | Helium level fluctuation on current lead B, leading to intermittent loss of CRYO OK, resulting in automatic SPA in the 14KA main power supply. The automatic SPA occurred twice (at 11400 A and at 12000 A after 11 min). ACR informed. They asked to wait for 20 min and then try again. | He consulted ACR team and found out that due to fluctuation in the temp of current lead B the cryo ok is lost while the test is in progress. ACR team will watch the temp. while the test is carried out next time. |
| “ | Third trial also failed. The magnet stayed at 12000 only for 5 minutes. Frederico was called. |  |
| TBE2 | Pierre told that the HV test at warm will have to be done at safe voltage levels, requiring new sequence to be defined and run. Therefore he insisted that Frederico will have to do this test. | Frederico has done HV(warm) & reported that it has failed and asked to kept as it is and told that he will investigate tomorrow. |
| A,B,C | Cluster A,B,C : Magnet name missing in Tema. No tests could be done. | Frederico called Alessandro Reimondo. He came and set right the problem |
| TBB1 | Problem in Quadrupole Vs. ground step of HV test (CDBP). TC contacted. He in turn talked to Pierre. Pierre instructed OP team to carry out provoked quench at 1.5 KA and then repeat the HV test. The HV test, after provoke quench, have failed again. |  |

| 02N(WED) |
| --- |
| TBB1 | On B1, PT2 failed at Quad Vs Gnd after 3kA Quench. |  |

| 03M(THU) |
| --- |
| TBE2 | HV done manually by ES at 1000V with power cables removed: result is ok. HV @ warm to be repeated as soon as the power cables are reconnected this afternoon. |  |
| TBF2 | Prep 5: Initial convalescence voltage was observed to be 3.17V instead of ~11.5V. This is not the case when P10 cable is removed | Frederico investigating the cable fault & cleared |

| 03A(THU) |
| --- |
| TBD2 | New magnet S4-618 has mismatch between bar code label, alphanumeric label and name in MTF database. Bench configured with MTF database name |  |
| TBE2 | Power cables connected. Warm HV (Prep 3) done as instructed by eqpt.support. |  |
| “ | HV result OK. Frederico suggested to cool down the magnet again and launch cold HV. TC also recommended to cool down the magnet, and to consider it as not repeated one. |  |

| 03N(THU) |
| --- |
| TBB1 | Eqpt. support contacted for HV problem on TBB1. Prep3 (HV at warm) was done Cryo was informed to cool down the magnet Since the E-Traveller was already signed till sL. No. 7.3. The cryo operator requested to un-sign the e traveller for cooling down. Hence two last signatures were unsigned. Further un-signing is not allowed by the E- Traveller. Cryo team is informed. | Problems with going back in e-traveller |

| 04M(FRI) |
| --- |
| TBD2 | Anticryostat end-plate yet to be connected on aperture 1. ICS has signed for warm tests. Complete scanning not possible as end plates are not put in place on the MRB side. IFS box seal not opened. Contacted ICS and asked them to make the IFS box accessible for carrying out warm tests. Halik informed that the connections are completed and electrical tests can continue. | J Halik informed that some connections are yet to be made and then the end plates will be put. It was done later. |

| 04A(FRI) |
| --- |
| TBB1 | After warming up and cooling down, CQBP again failing at Quderpole Vs Gnd second step 900 Volt as voltage > 500 Volt. Bypassed this step and launched the test again, it cleared rest of all the steps. So inform to equipment support, they suggested to do 3.0 KA quench. But if the HV fails warm up. |  |
| 04N(FRI) |
| TBA1 | During PTC 6 instability in the cryo occured at current 12850A after 3 min hold, which resulted in automatic SPA trigger & current ramp down to zero. Informed to TC & TC suggested to repeat the test again. |  |
| TBB1 | (SSS 308): PT 2 again failed Quad Vs Gnd failed at 900V step. As per information given by the last shift magnet is put for warm up. |  |

| 05M(SAT) |
| --- |
| TBE2 | for 2281:- Cold HV (CDBP) failed again. TC instructed to perform a provoke quench @1.5kA (all HF) and repeat HV manual step. If PT2 fails again magnet can be warmed up. |  |
| TBB1 | for 308:- Cold HV (CDBP) failed after 2nd attempt of cooldown. Tc asked to warm it up . |  |

| 06A(SUN) |
| --- |
| TBB2 | Unable to launge the correctors tests CD & EF => PC Aux fault / ES and PO piquet called : PC fault or /and regulation problem |  |
| “ | With the PO piquet, PLCSC was reset a 2ndtime.'PC Aux' error disepeared. | Done on 7th morning. |

07M(MON) Water supply problem : No cold tests could be done

| 07A(MON) |
| --- |
| TBE2 | 2281: TC (M Buzio) said that he will ask the ICS to proceed with disconnection of magnet so that repair work can be done off the bench. |  |