### EV/Week 35 starting 28 aug 06

### SM18 Operation Meeting (Tuesdays) 29 Aug 06

Bench status  on Monday  28-08-2006 at 8/28/2006 1:57:01 PM

| 5 | A1 | MBAL1403 | Prep 0 Scan Bench | 296.79 |
| --- | --- | --- | --- | --- |
| 2 | A2 | MBBR2316 | PT 12 Warm Up | 207.52 |
| 6 | B1 | SSS311 | PT 8 Fast De-excitation | 1.90 |
| 7 | B2 | SSS501 | Prep 0 Scan Bench | 296.88 |
| 8 | C1 | SSS296 | PT 14 Magnet MAPS | 295.8 |
| 9 | C2 | SSS241 | Prep 0 Scan Bench | 296.34 |
| 3 | D1 | SSS662 | PT 13.2 Resist. Meas. | 296.7 |
| 1 | D2 | SSS667 | Finished Ready to go Final scan | 296.43 |
| 10 | E1 | MBBR2369 | PT 5.2 1.5 KA Quench HF | 1.90 |
| 4 | E2 | MBBR2332 | Prep 5 Cool Down | 292.85 |
| 11 | F1 | MBAL1404 | PT 12 Warm Up | 203.71 |
| 12 | F2 | MBBR1346 | Cryo problem | 296.25 |

Tempr. from TBSM18 site

CONFIG :  D 6 - AS 3 -  IRSSS 3 ;

Supervisor:  Vittorio  ;Eq support : Richard 16-3602

14 Magnet Tests completed (  )

7 Dipoles “STAND BY” for : 	2274 (more than 9 ramp during for taining)

1367 (HV problem)

4 Arc-SSS  tests (318,313,345,140)

1  500-series [ 510 OK, 1 time mechanical problem ]

2  600-series [ 660 (148 h), 625 (144 h) ]

General Points:

- News /Situation SSS501 from  :  Mechanics team

  - B2 : SSS501 magnet came on bench with” line N” close. CFB will be modified to accept specific magnet with line N not conform.

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

  - F2 : MB 1346,  problem solved  Monday17:00 cooling down

  - D1 : SSS662 He leak cold mass.

- Situation SSSS  :D1 & D2 issues;

  - D1 : SSS662 magnet with leak problem under disconnection/remove;

  - D2 : New magnet arrived in SM18, under preparation for testing.

- A.O.B..

WEEK 34

Benches A, C, E, F

Bench D

Bench B

# From 8/20/2006 to 8/26/2006

Tests Information

| Magnets under Test from 20-08-2006 to 26-08-2006 |
| --- |
| - | Bench | Magnet | Shafts | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |
| 1 | TBA1 | MBAR2371 |  | 17-08-2006 | 30 h. | 26 h. | 19 h. | 16 h. | 20 h. | 22-08-2006 | 111 h. | Stripping |
| 2 | TBA2 | MBAR1322 |  | 19-08-2006 | 24 h. | 30 h. | 18 h. | 22 h. | 32 h. | 24-08-2006 | 126 h. | Stripping |
| 3 | TBB1 | SSS318 |  | 16-08-2006 | 33 h. | 12 h. | 128 h. | 8 h. | 50 h. | 26-08-2006 | 231 h. | Stripping |
| 4 | TBB2 | SSS510(2) |  | 17-08-2006 | 23 h. | 13 h. | 44 h. | 12 h. | 17 h. | 22-08-2006 | 109 h. | Stripping |
| 5 | TBC1 | SSS313 |  | 14-08-2006 | 27 h. | 28 h. | 33 h. | 7 h. | 33 h. | 20-08-2006 | 128 h. | Stripping |
| 6 | TBC1 | SSS345 |  | 20-08-2006 | 30 h. | 17 h. | 17 h. | 8 h. | 46 h. | 25-08-2006 | 118 h. | Stripping |
| 7 | TBC2 | SSS140 |  | 18-08-2006 | 23 h. | 15 h. | 22 h. | 7 h. | 15 h. | 22-08-2006 | 82 h. | Stripping |
| 8 | TBD1 | SSS660 |  | 17-08-2006 | 42 h. | 15 h. | 65 h. | 10 h. | 16 h. | 23-08-2006 | 148 h. | Stand By |
| 9 | TBD2 | SSS625 |  | 18-08-2006 | 36 h. | 23 h. | 53 h. | 9 h. | 23 h. | 24-08-2006 | 144 h. | Stand By |
| 10 | TBE1 | MBAR1302 |  | 16-08-2006 | 22 h. | 21 h. | 28 h. | 15 h. | 12 h. | 20-08-2006 | 98 h. | Stripping |
| 11 | TBE1 | MBBR1367 |  | 20-08-2006 | 30 h. | 22 h. | 10 h. | 15 h. | 48 h. | 26-08-2006 | 125 h. | Stand By |
| 12 | TBE2 | MBBR2359 |  | 19-08-2006 | 21 h. | 30 h. | 15 h. | 14 h. | 11 h. | 23-08-2006 | 91 h. | Stripping |
| 13 | TBF1 | MBAR2370 |  | 17-08-2006 | 29 h. | 37 h. | 72 h. | 14 h. | 33 h. | 25-08-2006 | 185 h. | Stripping |
| 14 | TBF2 | MBAR2274 |  | 07-08-2006 | 30 h. | 56 h. | 268 h. | 16 h. | 13 h. | 23-08-2006 | 383 h. | Stand By |

Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR2371 | 11137.9 A. (D2_L) | 12242 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - |
| 2 | MBAR1322 | 11354.8 A. (D1-U) | 12193.7 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - |
| 3 | SSS318 | 11583.7 A. (QE_34) | 11768.4 A. (QE_12) | 12353.7 A. (QE_34) | 12527.9 A. (QI_12) | 12610 A. (QE-34) | 12651.7 A. (QE_12) |  | - | - | - |
| 4 | SSS510 | 11656.5 A. (QI_12) | 11881.7 A. (QE_34) | 12389.4 A. (QE_12) | 12598.5 A. (QE_34) |  | - | - | - | - | - |
| 5 | SSS313 | 11464.7 A. (QE-34) | 12117.7 A. (QE-12) | 12409.5 A. (QE_12) |  | - | - | - | - | - | - |
| 6 | SSS345 | 12018.5 A. (QI_34) | 12426 A. (QE_34) | 12570 A. (QE_12) |  | - | - | - | - | - | - |
| 7 | SSS140 | 11232.2 A. (QI_12) | 11900.5 A. (QI-34) | 12319.9 A. (QI-34) | 12512 A. (QE-12) |  | - | - | - | - | - |
| 8 | SSS660 | 3211 A. (MQY-EXT) | 3655 A. (MQY-EXT) | 0 A. (No Data) | 3761.4 A. (I-MQY_34) | 3788 A. (I-MQY-12) | 3826 A. (I_MQY_34) | 3869 A. (i_mqy_12) | 3852.8 A. (In_QY_12) | 0 A. (No Data) | 3847.5 A. (ExtMQY34) | 3896.8 A. (IntMQY12) | 3900 A. |
| 9 | SSS625 | 5293.1 A. (Ex_MQM_2) | 5657 A. (Ex_MQM_2) | 5820 A. | - | - | - | - | - | - | - |  |  |
| 10 | MBAR1302 | 12121.5 A. (D2_U) | 12807.3 A. (D1_U) | 12239 A. (D1_U) | 12337.7 A. (D1_U) | 12850 A. | - | - | - | - | - |  |  |
| 11 | MBBR1367 | 11785.2 A. (D1_U) | 12847.1 A. | - | - | - | - | - | - | - | - |  |  |
| 12 | MBBR2359 | 12284.6 A. (D2_L) | 12857.1 A. | - | - | - | - | - | - | - | - |  |  |
| 13 | MBAR2370 | 11503 A. (D1_U) | 12156 A. (D1_U) | 12761.5 A. (D2_L) |  | - | - | - | - | - | - |  |  |
| 14 | MBAR2274 | 11498 A. (D2_L) | 11598.4 A. (D1_L) | 11538.4 A. (D1_L) | 11984.2 A. (D1_L) | 12286.8 A. (D1_L) | 12529.9 A. (D1_L) | 12646.2 A. (D1_L) | 12562.4 A. (D1_L) | 12694.8 A. (D2_L) |  |  |  |

2274 : MM + Not reach ULTIMATE.

SSS625/660 tested fast…