### VC/Week 37 starting 11sept 06

### SM18 Operation Meeting (Tuesdays) 12 sept 06

Bench status  on Monday  11-09-2006 at 9/11/2006 9:08:58 AM

| 8 | A1 | MBAR2274 | Prep 5 Cool Down | 99.93 | Mistake not have done ROBTC |  | Last time |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 4 | A2 | MBAR1338 | PT 12 Warm Up | 13.25 |  |  |  |
| 5 | B1 | SSS329 | PT 6.1 Training 1 | 1.93 |  |  |  |
| 1 | B2 | SSS309 | PT 13.2 Resist. Meas. | 298.58 |  |  |  |
| 7 | C1 | SSS298 | Prep2.2C (CD clamp) | 297.41 | HV fail- not trained |  |  |
| 3 | C2 | SSS339 | PT 6.6 Training 6 | 4.51 HV |  |  |  |
| 2 | D1 | SSS526 | PT 3 QH Measure | 1.91 | ROBTC |  |  |
| 10 | D2 | SSS608 | Prep 3 CDW1 HV | 298.76 | YT318 no antC |  |  |
| 9 | E1 | MBAR1374 | Prep 5 Cool Down | 262.41 |  |  |  |
| 11 | E2 | MBAL1401 | ICS 3 Traveller after warm test | 299.2 |  |  |  |
| 6 | F1 | MBAL2168 | PT 13.2 Resist. Meas. | 298.58 | For cryo leak |  |  |
| 12 | F2 |  |  | 298.58 |  |  |  |

CONFIG :  D 6 - AS 3 -  IRSSS 3 ;

Supervisor:  vittorio  ;Eq support : federico 16-3602

17 Magnet Tests completed (  )

9 Dipoles tests  ( 1369 ROBTC)

6 Arc-SSS  tests ( 309 put tro ROBTC .. rest stripped)

1  500-series [ 501 (193 hr) ]

1  600-series [  Q6 680 (214 h) , 6 MQTLs  ]

General Points:

- News /Situation  : D1 608 retest  for YT318. No Anticryo. Faster Training now ?

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- Situation SSSS  :D1 & D2 issues;

- A.O.B..

  - Problems last Thurs/fri on B IAP at warm /disconnect/check/  : Federico

  - D1 D2 switching issues wed & problems (Giorgio/salem)

  - + later on Federeico corr powering on D

| Magnets under Test from 03-09-2006 to 09-09-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR2379 |  |  | 02-09-2006 | 35 h. | 41 h. | 25 h. | 18 h. | 32 h. | 09-09-2006 | 151 h. | Stripping |  |  |
| 2 | TBA2 | MBBR1369 |  |  | 31-08-2006 | 28 h. | 19 h. | 35 h. | 18 h. | 0 h. | 04-09-2006 | 100 h. | Not Found or Thermal cycle |  |  |
| 3 | TBA2 | MBBR1369 |  |  | 04-09-2006 | 1 h. | 19 h. | 12 h. | 29 h. | 11 h. | 07-09-2006 | 72 h. | Stripping |  |  |
| 4 | TBB1 | SSS305 |  |  | 31-08-2006 | 25 h. | 11 h. | 35 h. | 8 h. | 23 h. | 04-09-2006 | 102 h. | Stripping |  |  |
| 5 | TBB2 | SSS309 |  |  | 05-09-2006 | 22 h. | 12 h. | 47 h. | 5 h. | 0 h. | 08-09-2006 | 86 h. | Not Found or Thermal cycle |  |  |
| 6 | TBB2 | SSS310 |  |  | 30-08-2006 | 49 h. | 13 h. | 26 h. | 7 h. | 39 h. | 04-09-2006 | 134 h. | Stripping |  |  |
| 7 | TBC1 | SSS315 |  |  | 01-09-2006 | 79 h. | 24 h. | 53 h. | 8 h. | 24 h. | 09-09-2006 | 188 h. | Stripping |  |  |
| 8 | TBC2 | SSS241 |  |  | 28-08-2006 | 34 h. | 15 h. | 16 h. | 6 h. | 77 h. | 03-09-2006 | 148 h. | Stripping |  |  |
| 9 | TBC2 | SSS346 |  |  | 03-09-2006 | 22 h. | 29 h. | 33 h. | 10 h. | 9 h. | 08-09-2006 | 103 h. | Stripping |  |  |
| 10 | TBD1 | SSS501 |  |  | 27-08-2006 | 87 h. | 17 h. | 28 h. | 9 h. | 52 h. | 04-09-2006 | 193 h. | Stand By |  |  |
| 11 | TBD2 | SSS680 |  |  | 30-08-2006 | 35 h. | 15 h. | 101 h. | 13 h. | 50 h. | 08-09-2006 | 214 h. | Stand By |  |  |
| 12 | TBE1 | MBAL1386 |  |  | 04-09-2006 | 36 h. | 27 h. | 18 h. | 15 h. | 27 h. | 09-09-2006 | 123 h. | Stripping |  |  |
| 13 | TBE1 | MBAL1399 |  |  | 30-08-2006 | 30 h. | 23 h. | 16 h. | 15 h. | 17 h. | 04-09-2006 | 101 h. | Stripping |  |  |
| 14 | TBE2 | MBAR2376 |  |  | 31-08-2006 | 45 h. | 21 h. | 13 h. | 14 h. | 27 h. | 05-09-2006 | 120 h. | Stand By |  |  |
| 15 | TBF1 | MBAL1383 |  |  | 29-08-2006 | 31 h. | 24 h. | 19 h. | 14 h. | 34 h. | 03-09-2006 | 122 h. | Stripping |  |  |
| 16 | TBF1 | MBBR2360 |  |  | 03-09-2006 | 24 h. | 22 h. | 24 h. | 16 h. | 8 h. | 07-09-2006 | 94 h. | Stripping |  |  |
| 17 | TBF2 | MBBR1343 |  |  | 01-09-2006 | 27 h. | 21 h. | 14 h. | 15 h. | 11 h. | 05-09-2006 | 88 h. | Stripping |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR2379 | 11369 A. (D1_L) | 11862.5 A. (D1_L) | 12600 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 2 | MBBR1369 | 11605.7 A. (D2_U) | 12272.1 A. (D2_L) | 11998 A. (D1_L) | 12241.5 A. (D1_L) | 12449.5 A. (D2_U) | 12360 A. (D2-U) | 12641 A. (D2-U) | 12634 A. (D2-U) | 12722 A. (D2_U) |  | - | - | - |
| 3 | MBBR1369 | 12470.4 A. (D1_L) | 12723 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS305 | 12120.9 A. (QI_12) | 12238 A. (QE_34) | 12311.2 A. (QE_34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS309 | 12430 A. (QE-34) | 12591 A. (QE_12) | 12824.5 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS310 | 9837 A. (VClamp E) | 12735.9 A. (QE_34) | 12801 A. (QE-12) | 11241 A. (VClamp I) | 11241 A. (VClamp I) | 12816 A. (QE_12) |  | - | - | - | - | - | - |
| 7 | SSS315 | 9225 A. (QE-34) | 12137.8 A. (QI_12) | 12313 A. (QI_34) |  | - | - | - | - | - | - | - | - | - |
| 8 | SSS241 | 11169.3 A. (QE_34) | 12207.2 A. (QE_12) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 9 | SSS346 | 11770.9 A. (QI_34) | 11938.5 A. (QE_34) | 12179.7 A. (QE_34) | 12350.6 A. (QE_34) | 12577.5 A. (QE_12) | 12715.8 A. (QI_34) | 12827 A. (QE_12) | 12850 A. | - | - | - | - | - |
| 10 | SSS501 | 12300.1 A. (QE_12) | 12424.1 A. (QE_12) | 12656.7 A. (QI_12) |  | - | - | - | - | - | - | - | - | - |
| 11 | SSS680 | 316 A. (No Data) | 395 A. (No Data) | 433 A. (No Data) | 428 A. (No Data) | 443 A. (No Data) | 467 A. (No Data) | 347 A. (No Data) | 472 A. (No Data) | 478 A. (No Data) | 474 A. (No Data) | 477 A. (No Data) | 483 A. (No Data) | 384 A. (No Data) | 412 A. (No Data) | 369 A. (No Data) | 484 A. (No Data) | 475 A. (No Data) | 0 A. (No Data) | 248 A. (No Data) | 296 A. (No Data) | 475 A. (No Data) | 490 A. (No Data) | 473 A. (No Data) |  |
| 12 | MBAL1386 | 11342.4 A. (D2_U) | 12395.8 A. (D1_L) | 12700.8 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |  |  |  |  |
| 13 | MBAL1399 | 11959.1 A. (D2_L) | 12556.5 A. (D1_L) | 12704.3 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |  |  |  |  |
| 14 | MBAR2376 | 11883 A. (D1-U) | 12377.8 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |  |  |  |  |
| 15 | MBAL1383 | 11083.3 A. (D2-U) | 12466.7 A. (D1-L) | 12779.9 A. (D1_U) |  | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |  |  |  |  |
| 16 | MBBR2360 | 11331.7 A. (D1_U) | 12466.7 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |  |  |  |  |
| 17 | MBBR1343 | 10917.1 A. (D1_L) | 12117 A. (D2_L) | 12758.8 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |  |  |  |  |  |  |  |  |  |  |  |

Long Training + extra tests due to anticryostat on 680 with 6 MQTL’s   .. also additional questionable tests with antic sealed

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_05_A | TBD2 | SSS680: Message received "Anticryostat Temperature too low". AC cable not connected. | MTM |  | Guy informed. He suggested proceeding ignoring the message. |
| Tuesday_05_N | TBD1 | SSS526: Clamp correctors check could not be done Permanent triggers due to the existing settings done by Giorgio to make the electronics compatible to the magnet s4 - 680. s. | MTM | IAP | Contacted Guy. He is not aware of the changes made by Giorgio. Awaiting intervention of Giorgio/mathias |
| Tuesday_05_N | TBC2 | SSS-346: Corrector powering on IJ could not be carried out as there is a note on the Power converter (IJ) instructing not to use the power supply. | AB/PO |  | IJ training will be carried out after the converted is rectified and carrying on with training. In fact this PC is no longer used for corrector test but the adjacent one. Richard comments: I have put some stickers on the rack with the spare 60Amps power converter. The ones we are using, to power cluster C, are the ones with the switch on top of them. It has always been like that. |
| Tuesday_05_N | TBB1 | SSS329 : Labels missing (for 1 and 2) on voltage tap connection leads. | CRI |  | ICS had problem to perform the connection. |
| Tuesday_05_N | TBC2 | Multimeter is not displaying the last digit. |  |  | The multimeter was set in fast sampling instead of medium sampling mode |
| Wednesday_06_A | TBD2 | SSS680 : Final HV is done and sign out E-traveller. Paper traveller pages not available. | MTM |  |  |
| Wednesday_06_N | TBD1 | SSS5326: IAP test in fault Vano and V cathode. | MTM |  | Federico informed. He asked to proceed. |
| Wednesday_06_N | TBE1 | MB1386: HV test at cold BP failed: dipole_quench-heaters vs ground. | MTM |  | Federico came and check that all the circuits are OK separately. He advised to perform PT3, PT4 and PT5 and then check again. |
| Wednesday_06_N | TBB2 TBC2 | SSS309 and SSS346: TCL_aux not OK | ACR |  | ACR team informed. Team leater advised to change the magnet priority to get the condition in TBB2. |
| Thursday_07_M | TBA2 | MB1369: Magnet held up at 290 K for7 hours. | MTM |  | 4 magnets warming up and MB1369 in low priority. Priorities changed. |
| Thursday_07_M | TBB2 | SSS309: Again training quench in the clamp. Still problem of flow in current leads | ACR |  | Cryo team takse the control in manuel mode. |
| Thursday_07_M | TBA2 | MB1369: IAP@warm. Offset error on VT_Cathode signal (Channel labelled VT_Anode on the potaim card). | MTM |  | Federico asked to try with another card. OK. Federico changed the faulty card. |
| Thursday_07_A | TBC1 TBB2 | SSS315: It takes a long time too cool down to 1.9 K SSS309: Difficulty to stabilize He flow in current leads | ACR |  |  |
| Thursday_07_A | TBB2 | SSS309: During PT8 quench is occurred in V-Clamp Ext near 12 KA current. | ACR |  | Federico informed to re-launch the test with increased Helium flow and during current ramping observe the V-Clamp voltage. Cryo team asked to OP team to check if condensation is on the top of CFB. OP team put a blower. The MQ quenched again in V_clamp after 3 mins. Cryo stability reached in the night. PT8 performed. |
| Thursday_07_N | TBB2 | SSS309: HV CDAP failed due to the condensation in the CFB (high level flow) despite the blower. All steps were ok except Quadrapole Vs M(X)D-M(X)F correctors & M(X)D-M(X)F correctors Vs Gnd failed. | ACR |  | PT9 executed manually. Federico asked to open a NC (failure at 1.5 kV) |
| Thursday_07_N | TBC1 | SSS315: HV (PT9) can not be launched due to LOT of Condensation on CFB. Humidity count = 68%.. | ACR? MTM? |  | Cryo team informed. All the parameters are OK from their side |
| Friday_08_M | TBB1 | SSS329: Cryo instability around 1.9 K | ACR |  | Cryo tema tried to solve the problem Action ACR : pump at warm in the “decanteur” to remove eventual humidity. |
| Friday_08_A | TBB1 TBD1 | SSS329, SSS526 : Triggers during IAP at cold “ Vclamp_ext and V_supra” | MTM |  | Federico brigged the Vtaps faulty. |
| Friday_08_A | TBF2 | MB1400 : TRU1 had a pin missing in the connector (ap2). | MTM |  | ICS changed TRU1 with TRU5. |
| Friday_08_A Saturday_09_A Saturday_09_A | TBB1 | SSS329: Quench during the 1kA ramp due to VClamp.( Vtaps broken) | ICS |  | Called Federico, TC and A.Siemko for advice. A.Siemko advised to warm_up the magnet and ICS will repeat the connection. Warm tests has to be performed before re-opening. Federico will inspect the V_taps during the opening. Manchette re-open: A.Siemko and Federico found an inversion between two V_taps and the wires were not soldered carefully. ICS has redone the connection. |
| Friday_08_N | TBD1 | SSS526: Corrector cannot be performed. During powering of CD/EF circuits the current ramp down at 13 A | MTM |  | Federico advised to train before the MQ. He will take action tomorrow. Federico and Siemko cleared a faulty status on the energy extraction system of the correctors. Federico called Sunday_morning. Extraction switch re-open Monday Giorgio closed the switch and un validated the command |
| Saturday_09_M | TBD2 | SSS608: HWREC is off. TEMA read the previous magnet name SSS680. | AB/CO | HWREC | Federico will take action. |
| Saturday_09_A | TBB2 | SSS309: HV CQBP OK after thermal cycle. | MTM |  | A.Siemko contacted for further instructions. He advised to continue the training with 2 quenches and warm_up with a training quench. |
| Saturday_09_A | TBF2 | MB1400: Warnings during MM in the angle differences (between backward and forward rotation) | MTM |  | Federico could not touch Stephane by phone. He advised to proceed. |
| Sunday_10_M | TBA1 | MB2274: Alarm message from ACR “temperature of anticryostats too low” |  |  |  |
| Sunday_10_M | TBA2 | MB1338: Quench at 1.5 kA for hf heaters: value of time discharge smaller in the graph than observed during the quench. Same results but attenuated. | AB/CO | AQA | Federico suggested trying again with manuel trigger.of QH card 2. Federico suggested informing TC for decision. |
| Sunday_10_A | TBB | The PC converter switch did not work : no light glowing | AB/PO |  | Federico asked to do it manually. OP TEAM managed to put the switch in CC after many trials |
| Sunday_10_A | TBC2 | SSS339: AQA analysis for 1st training could not be done as there was a buffer problem in DAQ server |  |  | Contacted Federico. He suggested generating file by manually triggering after DAQ arming. |
| Monday_11_M | TBD1 | SSS526: IAP at cold (PT-1) was failing and there were offset error on QE_34-12,Q_I_E_lg, Vgas_Ext_lead | MTM | IAP | Offsets measured in wrong outputs. Federico solved the problem. |