### VC/Week 39 starting 25 sept 06

### SM18 Operation Meeting (Tuesdays) 26 sept 06

Bench status  on Monday  25-09-2006 at 9/25/2006 9:29:03 AM

| 10 | A1 | MBBR2380 | PT 13.2 Resist. Meas. | 298.31 |
| --- | --- | --- | --- | --- |
| 4 | A2 | MBBR1370 | PT1 IAP@1.9K, Part III | 1.90 |
| 3 | B1 | SSS359 | PT 1 IAP @ cold | 1.90 |
| 11 | B2 | SSS139 | PT 13.2 Resist. Meas. | 298.04 |
| 9 | C1 | SSS144 | ICS 4 Final connection | 296.43 |
| 5 | C2 | SSS236 | Prep 5 Cool Down | 87.77 |
| 2 | D1 | SSS661 | PT 6.6 Training 6 | 4.60 HV |
| 12 | D2 | SSS673 | PT 13.2 Resist. Meas. | 296.79 |
| 6 | E1 | MBAR1376 | Prep 5 Cool Down | 120.23 |
| 7 | E2 | MBBR2375 | ICS 4 Final connection | 297.32 |
| 1 | F1 | MBBL1379 | PT 12 Warm Up | 226.54 |
| 8 | F2 | MBBR2385 | ICS 3 Traveller after warm test | 296.88 |

CONFIG :  D 6 - AS 5 -  IRSSS 1 ;

Supervisor:  Gijs -Gilbert  ;Eq support : 16-3602

15 Magnet Tests completed ( 12 Magnet Stripped )

8 Dipoles tests  ( 1334 HV fail)

6Arc-SSS  tests (308 (HV fail) )

0  500-series [ 526 (192 hr) with thermal cycle]

1  600-series [608  (245 hrs)]

General Points:

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- Situation SSSS  :D1 & D2 issues;

- MM problems on B2;

- other MM issues/problems this week Sun night …

- A.O.B..

| Magnets under Test from 17-09-2006 to 23-09-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAR1334 |  |  | 14-09-2006 | 30 h. | 27 h. | 35 h. | 19 h. | 26 h. | 20-09-2006 | 137 h. | Stand By |  |  |
| 2 | TBA2 | MBAR1375 |  |  | 17-09-2006 | 28 h. | 29 h. | 14 h. | 26 h. | 10 h. | 22-09-2006 | 107 h. | Stripping |  |  |
| 3 | TBA2 | MBBR2377 |  |  | 12-09-2006 | 25 h. | 27 h. | 17 h. | 17 h. | 18 h. | 17-09-2006 | 104 h. | Stripping |  |  |
| 4 | TBB1 | SSS134 |  |  | 15-09-2006 | 19 h. | 14 h. | 18 h. | 7 h. | 7 h. | 18-09-2006 | 65 h. | Stripping |  |  |
| 5 | TBB1 | SSS308 |  |  | 18-09-2006 | 26 h. | 12 h. | 35 h. | 20 h. | 8 h. | 23-09-2006 | 101 h. | Stand By |  |  |
| 6 | TBC1 | SSS338 |  |  | 16-09-2006 | 21 h. | 14 h. | 22 h. | 8 h. | 26 h. | 20-09-2006 | 91 h. | Stripping |  |  |
| 7 | TBC2 | SSS137 |  |  | 19-09-2006 | 37 h. | 20 h. | 15 h. | 8 h. | 12 h. | 23-09-2006 | 92 h. | Stripping |  |  |
| 8 | TBC2 | SSS357 |  |  | 15-09-2006 | 35 h. | 13 h. | 15 h. | 7 h. | 15 h. | 19-09-2006 | 85 h. | Stripping |  |  |
| 9 | TBD1 | SSS302 |  |  | 13-09-2006 | 26 h. | 14 h. | 18 h. | 7 h. | 55 h. | 18-09-2006 | 120 h. | Stripping |  |  |
| 10 | TBD2 | SSS608 |  |  | 08-09-2006 | 108 h. | 20 h. | 39 h. | 16 h. | 62 h. | 18-09-2006 | 245 h. | Stand By |  |  |
| 11 | TBE1 | MBBL2346 |  |  | 13-09-2006 | 28 h. | 24 h. | 19 h. | 14 h. | 7 h. | 18-09-2006 | 92 h. | Stripping |  |  |
| 12 | TBE1 | MBBR2378 |  |  | 18-09-2006 | 26 h. | 21 h. | 25 h. | 16 h. | 31 h. | 23-09-2006 | 119 h. | Stripping |  |  |
| 13 | TBE2 | MBAL2352 |  |  | 14-09-2006 | 25 h. | 26 h. | 20 h. | 15 h. | 24 h. | 19-09-2006 | 110 h. | Stripping |  |  |
| 14 | TBF1 | MBAL2280 |  |  | 17-09-2006 | 26 h. | 25 h. | 30 h. | 16 h. | 13 h. | 22-09-2006 | 110 h. | Stripping |  |  |
| 15 | TBF1 | MBBR1331 |  |  | 11-09-2006 | 48 h. | 34 h. | 19 h. | 17 h. | 20 h. | 17-09-2006 | 138 h. | Stripping |  |  |
| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| 1 | MBAR1334 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 2 | MBAR1375 | 12360.5 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 3 | MBBR2377 | 11860.7 A. (D1_U) | 12691.3 A. (D1-U) | 12700.7 A. (D2-U) |  | - | - | - | - | - | - | - | - | - |
| 4 | SSS134 | 11760 A. (QI_12) | 12070 A. (QI - 12) | 12161 A. (QI_34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS308 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 6 | SSS338 | 11788.3 A. (QI_12) | 12798.7 A. (QE_34) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS137 | 12226.3 A. (QE_12) | 12365.6 A. (QI_12) | 12562 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 8 | SSS357 | 11432 A. (QI-12) | 11898 A. (QE--34) | 12307 A. (QI-12) | 12524.9 A. (QE-34) |  | - | - | - | - | - | - | - | - |
| 9 | SSS302 | 11778.3 A. (QI_12) | 12637.9 A. (QE_12) | 12691.5 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 10 | SSS608 | 3488 A. (E-MQy12) | 3657 A. (I-MQY12) | 3741.4 A. (I_MQY_1) | 3786.2 A. (E_MQY_1) | 3840.8 A. (I_MQY_2) | 3813.6 A. (E_MQY_2) | 3858.5 A. (E_MQY_1) | 3900 A. | - | - | - | - | - |
| 11 | MBBL2346 | 11242.6 A. (D2_U) | 11977.4 A. (D1_L) | 12454.8 A. (D1_U) | 12638.6 A. (D2_L) |  | - | - | - | - | - | - | - | - |
| 12 | MBBR2378 | 10628.6 A. (D2_U) | 11366.3 A. (D1_U) | 11270.2 A. (D2_U) | 12464.9 A. (D1_U) | 11621.2 A. (D1_U) | 12727.8 A. (D2_U) | 12794.5 A. (D1-L) | 12859 A. | - | - | - | - | - |
| 13 | MBAL2352 | 10137.8 A. (D1_U) | 12027.3 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 14 | MBAL2280 | 11082.1 A. (D1-U) | 11669.4 A. (D2-U) | 12612.1 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 15 | MBBR1331 | 11979.8 A. (D1_L) | 12458.5 A. (D1_L) | 12670 A. (D1_U) |  | - | - | - | - | - | - | - | - | - |

HV Failures on 1334 & 308;    608 usual issues why connection/ disconnection  so long

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_19_M | TBA1 | MB1334: Temperature too low for Anti Cryostat. | MTM /ACR |  | Cryo team informed. |
| Tuesday_19_A | TBF1 | MB2280: AQA does not update MIITS manual. MB2280 : During PT5.2 file is not getting transferred So unable to register MIITs values | MTM |  | Giorgio asked OP team to proceed and suggested to start the next test after turning LF and HF off. Giorgio contacted Two magnet folders were copied to vol12 via sm18op file manager and not from the one that Giorgio left opened. Then the owner of the directory is sm18op, therefore the automatic copy (incaa based) could not work. If it is needed to copy manually to vol12, do it from file manager as INCAA. |
| Tuesday_19_N | TBB2 | SSS139: Due to cryo problem (Cryo1_x2 is not ok), SSW for aperture 2 could not be launched. The roll angle test was launched. SPA in converter and simultaneously the wire broke. new wire installed | ACR MTM |  | Cryo team is informed. . Solved after a while. |
| Tuesday_19_N Wednesday_20_M | TBB2 | SSS139: TCLaux fault. SSS139: SSW on Aperture 2 Field strength at nominal was launched at about 8:30 after taking clearance form cryo but quench occur during run. It was also launched at 12:00 but current ramp down due to Cryo instability | ACR |  | Cryo team asked OP team to wait for 30 min. Cryo team was informed. Stéphane suggested to perform SSW Field strength at 5kA. Stéphane advised to stop the SSW @cold and continue with the MM due to shortage of time. |
| Tuesday_19_N | TBF1 | MB2280: PT6.1 couldn't be launched on F1 due to error- 222 and bad state 33 of Power Converter. The power converter was not getting ON. | MTM |  | Giorgio will be contacted in the morning. |
| Tuesday_19_N | TBB1 | SSS308: HV failed on Quadrupole Vs Ground step. Others are OK. Report is generated. | MTM |  | Giorgio will be contacted. Faults confirmed. IFS box removed and investigation by Giorgio and Richard. |
| Tuesday_19_M | TBB2 | SSS139: Temperature too low for Anti Cryostat. | MTM |  | AC cable disconnected for shaft introduction. |
| Wednesday_20_A | TBB2 | SSS139- MM1 error on current reading on Ap1 and Ap2 , error on Temperature reading . SSS139: Concerning temperature error, wrong bench was selected in magnet settings in MMP. This was first time MM on this bench. | MTM |  | Guy called. Current reading cable was not connected due to previous field advance measurement. Guy connected it, then OK. Guy corrected the magnet settings, then temperature value correct. |
| Wednesday_20_N | TBB2 | SSS139: Errors in TEMA settings for MM3. | MTM |  | Guy corrected the errors and launched the load-line. |
| Wednesday_20_N | TBB2 | SSS139 Quench in (+)Vclamp_ext during MM3 Full load line at 11850A. | ACR |  | Cryo contacted. Quench occurred in Vclamp twice. Cryo investigation: Condenser problem. It will be solved for the next shift. |
| Wednesday_20_N | TBC1 | SSS244: Error-2 while doing IAP with HF heater discharge. MXI crate reset and restart. The crates were switched off and put on and IAP was done again. The problem repeated. Therefore the crates were put off and restarted again. This time the error was reported after the file was generated. | AB/CO | IAP |  |
| Wednesday_20_N | TBC2 | SSS137: OK to cool-down came after 24 hrs of ICS3 step as cryo leak problem. But after 1 hrs of ok to cool down sign, cryo inform that they can not cool down as anticryostat problem. | ACR |  | Why? Is there anti-cryostats in SSS137? |
| Thursday_21_A | TBD1 | SSS661: Error in IAP at warm, HF discharge of QH2,QH4, tminV>1000 | MTM | IAP | Guy informed. He advised to proceed. |
| Thursday_21_A | TBF2 | MB1342: Leak test not conform. Connection checked at warm by ACR | ACR |  | Leak in a flexible inside the CFB. Repaired but magnet to be cooled down again to validate the cryogenic integrity. |
| Thursday_21_A | TBD1 | SSS661: Temperature of anti-cryostats too low. Alarm displayed | /MTM |  | AC cables not connected. |
| Thursday_21_A | TBB2 | SSS139: Corrector magnetic measurements. The 600A power supply triggered at -300 A. | AB/PO |  | AB/PO called. Problem in tuning the power supply parameters? |
| Thursday_21_N | TBB2 | SSS139: MM2 + ramp tests: Error in MM position at high current and low bucking ratio for 1 sector | MTM | MMP | Stéphane informed. Results checked. OK to proceed. |
| Friday_22_A | TBA1 | MB2380: Water inlet point on CFB (blue mark steel pipe) leaks. It sprays when attempt to connect the feed pipes. | MTM |  | The pipe is properly connected after good efforts but repair may be required. Giorgio triggered an action with the mechanic team |
| Friday_22_A | TBA1 | MB2380: Min Energy Quench at 3kA. Error displayed: "Generic error message window" and no LF data available while analysis. After resetting the crates , no success. | AB/CO | lf | Guy was informed. He asked to reset the crates |
| Friday_22_N | TBA1 | MB2380: PT6.1, message "ERROR, LF cannot be armed" appeared on TEMA display. It couldn't be solved even after resetting the LF & HF crates near test bench.. | AB/CO | Hf,lf | Guy was informed. He came and changed the fan unit of HF crate, reset the crates, as well as the sunmta31 |
| Friday_22_N | TBB2 | SSS139: HV test for Orbit dipole corrector. Quadrupole vs orbit dipole and quench heater vs ground failed as there was condensation on top of CFB. | AB/CO | Hf,lf | OP team cleaned water from voltage taps and blower was applied and the repeat steps were ok. |
| Sunday_24_N | TBB1 | SSS359: MXI error during PT1. DAQ (hf, lf) cannot be armed. | AB/CO | Hf,lf | Federico informed in the morning, he killed the process which was hung. |
| Monday_25_M | TBB1 | SSS359: TcL-aux fault. Corrector powering cannot be launched.. | ACR |  |  |